## Sheet1
LT13
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Dept : | Surgical Oncology (Sangrur, Punjab)
Name of equipment : ICU Ventilator - Qty- 03 nos.
Please quote all the models including the state of the art equipment which meets the basic minimum specifications.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model Name :
Make :
1 | Microprocessor controlled ventilator with integrated facility for ventilation suitable for pediatric to adult ventilation
2 | Graphic & numeric display with clear screen
3 | Facility to measure and display
a | atleast two waves-Pressure and time, Volume and time and flow and time waveforms
b | Flow volume and Pressure-volume loops. (Optional)
c | Status indicator for ventilator mode, battery life, Patient Data, Alarm setting, clock etc.
4 | Each unit be supported by compressor or turbine air supply with a its guarantee of 7 years (inclusive of all spares/consumables)
5 | Following settings for all age groups
a | Tidal volume
b | Peak insp. pressure
c | Pressure ramp or rise time
d | Respiratory rate
e | CPAP/PEEP
f | FiO2
g | Pause time
h | Pressure&/or flow trigger
6 | Monitoring of the following parameters
a | Airway pressure (Peak, Plateau & Mean)
b | Tidal Volume (Inspired and Expired)
c | Minute volume (Inspired and Expired)
d | Spontaneous Minute Volume
e | Total frequency
f | FiO2
g | Intrinsic PEEP
h | Resistance and compliance
i | Use selector Alarms for all measured & monitored parameters
7 | Modes of ventilation
a | Volume controlled or equivalent
b | Pressure controlled or equivalent
c | Assist controlled (Pressure and Volume)
d | Pressure support or Equivalent
e | SIMV (Pressure control & Volume control) with pressure support
f | CPAP/PEEP
g | Non invasive ventilation with leak compensation
h | APRV (Optional)
i | PRVC or equivalent (Optional)
8 | Apnoea/Backup ventilation
a | Flow measurement (Please specify)
b | Battery backup for minimum 30 minutes
c | Demonstration of the quoted model
9 | Accessories- Humidifier
a | Nebuliser to be used both off and On line
b | Reusable silicone autoclvable sets of each pediatric and adult hoses
10 | Warranty & After sales service
a | Should have local service facility
b | Must submit user list & performance report with in last 5 years from major hospital
c | Back to back warranty to be taken by the supplier from the principal to supply spares for minimum 10 years
11 | Expected life of equipment- please mention
12 | Warranty : 2 years
13 | CMC : 8 Years
1) | If waranty and comprehensive maintenance contract values are not mentioned TMC, reserves right to reject the offer.
2) | Please mention whether UPS/stabiliser is required for the equipment. If so should be in the offer.
3) | Only Yes or No should be mentioned in Compliance report and additional information
should in remarks column.