## Sheet2
HBCH/2021-22/CAP/PT/07
TATA MEMORIAL CENTRE
HOMI BHABHA CANCER HOSPITAL, SANGRUR
(PURCHASE DEPARTMENT)
Dept- Surgical Oncology
Technical Specification of LOWER GI QTY. 1 No. AND UPPER GI Qty- 01 No.
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of Installation at TMC
Sr. No. | Item description | Compliance |  | Remarks
A | UPPER GI ENDOSCOPE - Qty. 1 No. | Yes | No
Model name :
Make:
Year of introduction:
Year of Probable end of support:
1 | Should feature HDTV image quality which displays clear images with high precision & accuracy in colour tone
2 | Should be fully immersible in disinfectent solution for effective cleaning
3 | Should have water proof connector facility to minimize the risk of damage due to accidental immersion
4 | Insertion section should have air water nozzle, light guide lens, objective lens, instrument channel outlet.
5 | Field of view should be 140 degree
6 | Direction of view should be forward viewing
7 | Depth of field - 4-100mm for closer observation
8 | Distal end outer diameter should be in the range of 9.2mm-9.4mm
9 | Insertion tube outer diameter should be in the range of 9.2mm-9.4mm
10 | Instrument channel diameter - 2.5mm-3mm
11 | Minimum visible distance should be 3mm (Normal) from the distal end
12 | Angulation - UP 210 degree, Down 90 degree, Left-Right 100 degree
13 | Working length should be in the range of 1000mm - 1100mm
14 | Should be supplied with all the standard accessories such as channel cleaning brush, injection tube, air water channel cleaning adaptor, suction cleaning adaptor, biopsy valve, biopsy forceps, mouth guard, Biopsy forceps, Guide wire, Leakage tester, water tank, mouth piece, cytology brush
15 | Scope should have optical image enhancement technology (NBI/FICE/I-Scan OE)
B | LOWER GI ENDOSCOPE - Qty. 1 No.
Model name :
Make:
Year of introduction:
Year of Probable end of support:
1 | Should feature HDTV image quality which displays clear images with high precision & accuracy in colour tone
2 | Scope should have optical image enhancement technology (NBI/I-Scan OE/FICE)
3 | Should be fully immerssible in disinfectant solution
4 | Should have no need to attach water resistant cap
5 | Should have auxilary water jet for mucosal cleaning
6 | Field of view should be 140 degree
7 | Observation range should be 3-100mm
8 | Distal end outer diameter should be 12mm-12.8mm
9 | Angulation should be 180 degree for Up-Down and 160 Degree for Right - Left
10 | Instrument channel diameter should be 3.5mm-3.8mm
11 | Minimum visible distance should be 5mm (Normal) from the distal end
12 | Insertion section should have air water nozzle, light guide lens, objective lens, instrument channel outlet.
13 | Working length - 1600mm-1700mm
14 | Should be supplied with all the standard accessories such as channel cleaning brush, injection tube, air water channel cleaning adaptor, suction cleaning adaptor, biopsy valve, biopsy forceps, mouth guard, Biopsy forceps, Guide wire, Leakage tester, water tank, Mouth piece, cytology brush
C | VIDEO PROCESSOR WITH LIGHT SOURCE - Qty. 1 No.
Model name :
Make:
Year of introduction:
Year of Probable end of support:
1 | Should have Xenon 300 or equivalent LED light source
2 | Should have digital signal output as HD-SDI/DVI
3 | Should have white balance, color tone adjustment (RGB), Automatic gain control, contrast, noise reduction, Automatic Iris control, Freeze, Pre-freeze features
4 | Should be compact & ergonomically designed
5 | Should have Optical Image enhancement settings (NBI/FICE/Iscan-OE)
6 | Recording of both still & moving images equipped with one touch connection of scopes
7 | Should have USB slot for image recording
8 | Electronic zoom from 1X-2X
9 | Patient information should be able to be displayed such as Patient name, Patient ID, Sex, Age, Date of Birth
D | MEDICAL GRADE MONITOR - Qty. 1 No.
Model name :
Make:
Year of introduction:
Year of Probable end of support:
1 | Should be full HD LCD monitor 27" with high resolution 1920X1200
2 | Should have low power consumption
3 | Aspect ratio should be 16:10
4 | Should have inputs and outputs such as HD/SD, Video, Y/C, DVI
E | CO2 ENSUFFLATOR - Qty. 1 No.
Model name :
Make:
Year of introduction:
Year of Probable end of support:
1 | Should have controlled flow rate settings
2 | Should be supplied with water container, cylinder hose and tubings for appropriate usage
3 | Should be provided with extra low flow tube also
4 | Maximum Pressure feed should 65KPa
5 | Should give alarm in case of overpressure, also when the CO2 cylinder is about to get empty.
F | Should be provided along with Desktop with printer, Suction machine, Endoscope trolley, Recording software with required hardware, UPS, External HDD of 2TB
1 | Should perform yearly calibration/Validation and preventive maintenance (2nos) during warranty & if entered into AMC (includes unlimited breakdown calls)/CMC (which will include spare replacement consumable spares, breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/International standards (As per NABH norms). To submit SOPs for PMS at the time of installation/Commissioning
2 | Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
3 | Please specify list & cost of consumables/consumable Spares (i.e. Spares need to be replaced at regular intervals, may be quarterly/half yearly/yearly such as annual maintenance kit etc.) if any. (Quote Rate)
4 | Please specify pre installation requirements (electrical, HVAC, Compressed air (please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour); if any of the above is essential)
5 | Please specify footprint size & weight.
6 | Demo of th equoted model, will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
7 | Predispatch inspection at factory will be required if desired by authorities.
8 | Please visit the site to check whether quoted model can be accomodated in the available space.
9 | Should be CE/FDA/BIS/ICMED approved
10 | Please specify power consumption
11 | The quoted model should have minimum 10 installation in INDIA unless it is a new equipment.
12 | Unpacking and shifting the consignment to the installation site is to be included in the scope of supply. Bidder/Manufacturer/Authorized service provider should take responsibility to lift/shift the consignment from unloading site to the installation site. Unloading site shall be Stores Department, HBCH, Sangrur. If needed, bidder has to arrange for the labourers at no charge to HBCH, Sangrur. (Before submitting the quotation, bidders may visit HBCH Sangrur to know unloading and installation site.
13 | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
14 | Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
15 | Warranty-2 years
16 | CMC-5% of Ex-works cost per annum will be considered for all bidders for 8 years after warranty. All bought out items used in system should also be included in the warranty. CMC should include upgrades/update cost, No separate software license fee will be payable.
17 | Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted.
18 | Quoted company should be ISO certified.
1) If 2 years warranty, 8 years CMC and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject offer.
2) Please mention whether UPS / Stabilizer is required for the equipment. If so, should be included in the offer.
3) Only yes or no should be mentioned in the compliance report and additional info in remarks column.