## Sheet1
TMH/TMH/2021-22/CAP/LT/0113
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Department - Speech Therapy, TMH.
Technical specification of "Digital Double Channel Diagnostic Audiometer", Quantity 1 no.
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Item Description | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Year of Probable end of support:
Should have following features:
1 | Should be able to perform pure tone audiometry with air conduction, bone conduction and free field tests, Pediatric assesment, High Frequency audiometry.
2 | Should have wide band or narrow band masking & Insert receivers should be provided.
3 | Sould have facility to present stimulus inputs from CD or Mic
4 | Frequency range:
a | Air conduction: 125Hz to 8KHz
b | Bone conduction: 250Hz to 8KHz
c | Free field: 125Hz to 8KHz
d | High frequency range: 8KHz to 20KHz
5 | Speech test from 0 db to 120 db
6 | Should have talk back, talk over facility
7 | Should have special tests such as pure tone, tone decay, SISI, ABLB, stanger tests
8 | Distortion tone should be <0.1%
9 | Should have membrane keypad
10 | Should have backlit LCD/LED display with Hz, dB parameters display
11 | System Should be provided with required hardware and database software.
12 | System should have following accessories:
a | Single insert for masking
b | Talk over Microphone built-in or external microphone
c | Monitor Head Set
d | Built-in speaker or ext. earphone or speaker
e | Assistant head phone
f | Desktop microphone
g | Monitor speaker
h | Talk back microphone
i | Patient responder
j | IP 30 insert phone 10 Ohm
k | Sound field speakers SP 90
l | Path cords for all transducer and VGA cables for monitors for 2 room situation
m | Extra LCD/LED Monitor Screen for VRA 2nos for each audiometer min 19inch screen with height adjustment Wall mounting plates for VRA - VGA & HDMI switch box for controlling multiple monitors trough remote control
n | All-in-one Computer system (with Core i5 Intel processor 11th generation or better, 8GB DDR4 RAM with 1TB SSD/M2 & 1 TB external storage SSD, 4 GB Graphics card, minimum 15.6 inch screen size, genuine windows 10 with license)or higher model compatible with Audiometer & VRA system, all-in-one(printer, scanner, copier) color printer with Wi-Fi connectivity,,Ethernet:10/100/1G LAN, Connectivity: VGA,USB, HDMI, Built-in Wi- Fi, Bluetooth, Sound Card: Intel High Definition Audio 2.0
13 | Should have facility to connect with HIS system.
14 | Unit should have CE/ FDA/ ICMED/ BIS/ IEC standards.
15 | Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards.
16 | Should be supplied with standard make UPS with atleast 30 mins backup. (Optional)
17 | Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
18 | Please specify pre installation requirements (electrical, HVAC etc.)
19 | Please specify footprint size & weight.
20 | Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
21 | Predispatch inspection at factory will be required if desired by authorities.
22 | Please visit the site to check whether quoted model can be accomodated in the available space.
23 | Please specify power consumption.
24 | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
25 | Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
1 | Warranty : 2 years
2 | AMC : Maximum 2% per annum for 8 years on ex-works cost after warranty.
3 | Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted.
1. If 2 years warranty, AMC maximum 2% per annum for 8 years on Ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/ Stabilizer is required for the equipment. If so, should be included in the offer.
3. Only yes or no should be mentioned in the compilance report and additional info in remarks column.