## Sheet1
TMH/TMH/2021-22/CAP/LT/0114
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Department - Speech Therapy, TMH.
Technical specification of "Diagnostic Oto-Acoustic Emission (OAE) with Video-Otoscope", Quantity 1 no.
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Item Description | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Year of Probable end of support:
A | The Diagnostic Oto-Acoustic Emission (OAE) system should have following features:
1 | The system should have Quick Screen, Standard diagnostic recording, Customizable Analysis time and minimum analysis time should be less than 20 seconds
2 | It should have Click stimulus or configurable tone pulse, Configurable stimulus intensity
3 | System should perform In-the-ear calibration
4 | It should have Real-time stimulus and probe fit monitoring
5 | OAE system should be provided with required database software Computer Interface via USB and required hardware
6 | System should be provided with Permanent OAE probe
7 | Should be provided with Probe Cleaning Accessories
8 | Probe tip kit should be provided-3Nos
9 | Should have overlay function.
10 | DPOAE :
a | Frequency range:1.5KHz to 10KHz
b | Frequency step:1 Hz
c | Level:40 to 70 dB SPL
d | Level step:1 dB
e | A/D resolution:24 bit
f | SNR criteria: Minimum SNR ≤ 3 dB
g | DP criteria: Minimum DP Level, SNR, Residual Noise, DP Reliability
h | Information Display: DP gram & summary table
11 | TEOAE :
a | Frequency range:0.5KHz to 4 KHz
b | Frequency step:1 Hz
c | Level:40 to 80 dB PeSPL
d | Level step:1 dB
e | A/D resolution:24 bit
f | SNR criteria: Minimum SNR ≤ 5 dB
g | TE criteria: Minimum TE Level, SNR, Minimum Sweeps
h | Information Display: Graph & summary table
B | Video-otoscope
1 | Light Source : LED with Fiberoptic
2 | Tip Size should be less than 4 mm
3 | Focal Distance: > 10 mm
4 | Should view Tympanic membrane with 90° angle with a field view of +/-45°
5 | Option to capture both Image & Video
6 | Should be compatible with existing computer and windows software.
7 | Video otoscope should be integrated with existing computer and required software for viewing, saving & retrieving data
8 | Image Saving Format: jpeg(still)
9 | Compatible Mains & Cables, All size Speculums, probes & all other standard accessories to be provided with the equipment
Both the systems should have facility to connect with HIS system.
Unit should have CE/ FDA/ ICMED/ BIS/ IEC standards.
Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABL norms). To submit SOPs for PMS at the time of installation/commissioning.
Should be supplied with standard make UPS with atleast 30 mins backup. (Optional)
Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
Please specify pre installation requirements [electrical, HVAC, Compressed air ( please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour) ; if any of the above is essential]
Please specify footprint size & weight.
Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
Predispatch inspection at factory will be required if desired by authorities.
Please visit the site to check whether quoted model can be accomodated in the available space.
Please specify power consumption.
Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
Response Time: Should not be more than 06 hrs from lodging a breakdown complaint on toll free or by email.
Warranty : 2 years
AMC : Maximum 2% per annum for 8 years on ex-works cost after warranty
Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted.
1. If 2 years warranty, AMC maximum 2% per annum for 8 years on Ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/ Stabilizer is required for the equipment. If so, should be included in the offer.
3. Only yes or no should be mentioned in the compilance report and additional info in remarks column.