## Sheet1
TMH/TMC/2021-22/CAP/LT/0117
TATA MEMORIAL HOSPITAL
PURCHASE DEPARTMENT
Department - Transfusion Medicine
Technical specification of "Plasma Thawing Bath", Quantity 4 nos. (Qty. 2 nos. each for HBCHRC Vizag and HBCHRC Mullanpur).
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Name of the Model:
Make:
Year of Introduction:
Year of Probable end of support:
1 | Microprocessor based control circuit with digital display for set and actual temperature.
2 | Operational temperature should be 37°C with very high accuracy of around + / - 0.2 degree C
3 | Stainless steel interior with the capacity of holding approx. 12 plasma bags simultaneously per cycle
4 | Drain port for draining water bath within 3-5 minutes
5 | Should be supplied with suitable racks for holding bags in vertical position.
6 | Should incorporate circular and hardware to achieve set point rapidly for obtaining temperature uniformity.
7 | Cover to keep the unit covered when not in use.
8 | Constant water agitation to ensure uniform bath temperature
9 | Should meet national and international standards ISO/CE/FDA/BIS/ICMED approved.
10 | Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards (as per NABL norms). To submit SOPs for PMS at the time of installation / commissioning.
11 | Performance certificates of IQ,OQ,PQ tests to be submitted at the time of installation/commissioning;if desired by the end user (optional).
12 | Should be supplied with standard make UPS with atleast 30 mins backup (optional).
13 | Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
14 | Please specify list and cost of consumables/consumable spares (i.e spares need to be replaced at regular intervals, may be quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
15 | Please specify pre installation requirements (electrical, HVAC, Compressed air (Plaese specify required air pressure), water requirement (RO/ DI/ Distilled water with its pressuer, flow rate per hour); if any of the above is essential])
16 | Please specify footprint size & its weight.
17 | Demo of the quoted model will be required if so desired by the user, after the opening of technical bid and prior to the opening of financial bid. This is for technical evaluation.
18 | Pre dispatch inspection at factory will be required if desired by authorities.
19 | Please visit the site to check if the quoted model can be accomodated in the available space.
20 | Please specify power consumption.
21 | Local Service Support : Should have local office and service support/service engineer for attending the breakdown calls.
22 | Response Time : Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
1 | Warranty : 2 years
2 | AMC : Maximum 2% per annum for 8 years on ex-works cost after warranty.
3 | Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted.
1. If 2 years warranty, AMC maximum 2% per annum for 8 years on Ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3. Only Yes or No should be mentioned in compliance report and additional information in remarks column.