## Sheet1
TMH/MUL/2021-22/CAP/LT/0118
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Department - Speech Therapy, HBCHRC Mullanpur.
Technical specification of "Comprehensive Auditory Evoked Potential System", Quantity 1 no.
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Item Description | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Year of Probable end of support:
1 | System should be provided with all the required hardware, software & essential accessories.
2 | Should be able to be used :
* Standard Auditory modality including ECoghG (Electrocochleograph)
* Auditory brainstem response
* MLR (Middle Latency response)
* LLR (Late latency response)
* Chained stimuli module
* P300/MMN module
* Auditory steady state response
* OAE (DP and TE)- Oto Acoustic Emission
3 | Should be provided with following accessories:
a | Insert Earphones kit
* Intensity: 0-130 dB SPL
* Stimulus Type: Clicks, Tones 125Hz- 8KHz and CE-CHIRP
b | TDH 39 Headphones Kit
* Intensity: 0-120 dB SPL
* Stimulus Type : Clicks, Tones 125Hz-12KHz and CE-CHIRP
c | Bone Conductor Kit (B71/81)
* Intensity : 0-98 dB SPL
* Stimulus Type: Clicks, Tones 250Hz- 8 KHz and CE-CHIRP
d | Gold plated cup electrodes to be provided (1 set extra)
4 | Latest PC configuration and printer compatible with the main system should be provided
5 | Should have certification of International Electrotechnical Commision Standard - IEC/CE/FDA/BIS/ICMED Standards for medical electrical equipment.
6 | Should perform yearly calibration and preventive maintenance (2 nos.) during warranty & if entered into AMC (includes unlimited breakdown calls) / CMC (which will include spares replacement, consumable spares, breakdown calls) Testing & measuring equipments used should be traceable to SI units through National /international Standards (as per NABH norms).
7 | Should be supplied with standard make UPS with atleast 30 mins backup.
8 | Please specify list and cost of consumables/ consumable spares ( i.e spares need to be replaced at regular intervals, may be quarterly/half yearly/ yearly such as annual maintenance kit etc.) if any.
9 | Please specify pre installation requirements (electrical, HVAC etc.)
10 | Please specify footprint size & its weight.
11 | Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
12 | Predispatch inspection at factory will be required if desired by authorities.
13 | Please visit the site to check whether quoted model can be accomodated in the available space.
14 | Please specify power consumption.
15 | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
16 | Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
1 | Warranty : 2 years
2 | AMC : Maximum 2% per annum for 8 years on ex-works cost after warranty.
3 | Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted.
1. If 2 years warranty, AMC maximum 2% per annum for 8 years on Ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/ Stabilizer is required for the equipment. If so, should be included in the offer.
3. Only yes or no should be mentioned in the compilance report and additional info in remarks column.