## Sheet1
TMH/MUL/2021-22/CAP/LT/0041
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Department - Surgery, HBCHRC Mullanpur.
Technical specification of "Battery Operated Drill & Saw System", Quantity 01 no.
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 10 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Year of Probable end of support:
A | Battery reamer/drill:
1 | Speed of drill and reaming should be atleast 250 RPM (without attachement) or More
2 | Maximum torque should be atleast 15Nm (130 in/lbs).
3 | Should be light weight, not more than 2kgs
4 | Should be autoclavable.
5 | Should have dedicated forward and reverse switch with safe mode.
6 | Should be supplied with standard cleaning brush and consumables like oil and spray etc.
7 | Should be provided with reaming attachment, quick coupling for drill bit, chuck with key and quick coupling for K wire-1 each
B | Battery oscillator:
1 | Should have speed of maximum 12000 cycles/min
2 | Should be light weight, not more than 2 kgs.
3 | Should be autoclavable.
4 | Should have safe mode.
C | Battery:
1 | Battery of NiMH type or Lithium Ion type
2 | Should have maximum charging time of 70 mins.
3 | 180 deg opening of battery housing for easy insertion of battery.
D | Battery charger:
1 | should allow atleast 4 different batteries to be charged at the same time
2 | Condition of the battery should be easily assessed.
3 | charger should diagnose battery levels
E | Sterilization Kit
Should accommodate all handpieces, attachment and accessories for autoclave
Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABH norms)
Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
Please specify pre installation requirements (electrical, HVAC etc.)
Please specify footprint size & weight
Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
Pre dispatch inspection at the factory will be required if desired by the authorities.
Please visit the site to check whether quoted model can be accomodated in the available space.
Please specify power consumption.
The unit should be CE/FDA/BIS/ICMED approved.
Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
Warranty : 2 years
CMC : Maximum 5% per annum for 8 years on Ex-works cost after warranty
Back to back assurance to be taken by the supplier from the OEM to supply spares for minimum 10 years and to be submitted.
1) If 2 years warranty, CMC maximum 5% per annum for 8 years on Ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject offer.
2) Please mention whether UPS / Stabilizer is required for the equipment. If so, should be included in the offer.
3) Only yes or no should be mentioned in the compliance report and additional info in remarks column.