## Sheet1
TMH/VIZ/2021-22/CAP/LT/0122
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Department - Dental & Prosthetics Surgery, HBCHRC Vizag.
Technical specification of "Dental Scaler with Tips", Quantity 2 nos.
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Item Description | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Year of Probable end of support:
1 | System should work on piezo technology. Should eliminate dental calculus, bacterial plague and tea stains on the teeth quickly without hurting the gum or the enamel.
2 | Should not have heat generation in handpiece
3 | Should have titanium tip adapter
4 | should give around 30000 or better straight lined ultrasonic oscillation per second
5 | Output Power supply should be 3W to 20W as desired for work
6 | Preferred to have turbo mode for fast removal of hard calculus & tartar with power supply 15W to 40W or better
7 | Total Power supply : 115V to 230V
8 | Insert Rotation: 360 Degree
9 | Tip Movements : Ultrasonic straight oscillation
10 | Water supply : 1 to 5 bar
11 | Working frequency : 28KHz to 31KHz
12 | The unit should comply CE/FDA/BIS/ICMED
13 | Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABH norms). To submit SOPs for PMS at the time of installation/commissioning.
14 | Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
15 | Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
16 | Please specify pre installation requirements [electrical, HVAC, Compressed air ( please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour) ; if any of the above is essential]
17 | Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
18 | Please specify power consumption
19 | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
20 | Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
1 | Warranty : 2 years
2 | AMC : Maximum 2% per annum for 8 years on ex-works cost after warranty
3 | Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted.
1. If 2 years warranty, AMC maximum 2% per annum for 8 years on Ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/ Stabilizer is required for the equipment. If so, should be included in the offer.
3. Only yes or no should be mentioned in the compilance report and additional info in remarks column.