## Sheet1
TMH/VIZ/2021-22/CAP/LT/0124
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Department - Dental & Prosthetics Surgery, HBCHRC Vizag.
Technical specification of "Dental X Ray", Quantity 1 no.
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Item Description | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Year of Probable end of support:
Image X Ray System (Main Unit)
1 | Should have high frequency and constant potential (DC generator) which enables concentrated and very high - density radiation to be emitted
2 | Should be wall mounted Intra Oral X Ray unit and should be feasible to convert it into floor mounted as per the site availability and as per end user's requirement. Accessories required for the same should be provided along with.
3 | Should have tube voltage - 60 to 70 kVp
4 | Should have minimum range of exposure time - 0.02 to 3.2 seconds
5 | Should enable optimal diagnosis of intra-oral X-rays with a 35% reduction in exposure time
6 | The exposure time should be suitable for digital radiology with CCD, CMOS or phosphor plate sensors, compatible with RVG
7 | Generator should be compatible with all digital imaging systems
8 | Should have better patient protection
9 | Should have easy adjustable timer setting
10 | Should have excellent filtration in pure aluminium to protect the skin from soft radiation
11 | Should have leakage protection from very high radiation
12 | Should have very small focal spot diameter less than or equal to 0.7 mm for better sharpness
13 | Should have angle degree indicator on the head support for accurate vertical positioning
14 | Should have 395 degrees rotation of the head around horizontal axis
15 | Should have Microprocessor-controlled timer
16 | Should have Soft touch control panel
17 | Should have Digital display
18 | Should be error detector system
19 | Should be Class 1 type B model
20 | Should have 230VAC - single phase 50 / 60 Hz of voltage supply
21 | Should have Anode current of 4mA-8mA
22 | Should have Leakage radiation - Less than 0.25mGy/hour maximum at a distance of 1m from focal spot
23 | Should have X-ray emission control
24 | Should be AERB Approved
25 | Manufactured with international safety standards for radiation leakage
Standard Accessories
1 | Should be supplied with lead apron, gonaldal sheath, Thyroid protection collar
2 | Electrical cord for timer box with 12" cone, arm assembly with mains cord
General terms:
1 | Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABL norms). To submit SOPs for PMS at the time of installation/commissioning.
2 | Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
3 | Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
4 | Please specify pre installation requirements [electrical, HVAC, Compressed air ( please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour) ; if any of the above is essential]
5 | Please specify footprint size & weight
6 | Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
7 | Pre-dispatch inspection at factory will be required if desired by authorities
8 | Please visit the site to check whether the quoted model can be accomodated in the available space.
9 | Please specify power consumption
10 | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
11 | Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
1 | Warranty : 2 years
2 | CMC : Maximum 5% per annum for 8 years on ex-works cost after warranty
3 | Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted.
1. If 2 years warranty, CMC maximum 5% per annum for 8 years on Ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/ Stabilizer is required for the equipment. If so, should be included in the offer.
3. Only yes or no should be mentioned in the compilance report and additional info in remarks column.