## Sheet1
TMH/VIZ/2021-22/CAP/LT/0125
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Department - Surgery, HBCHRC Vizag.
Technical Specification of "Sternum Saw System", Quantity 2 nos.
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model:
Make:
Year of Introduction:
Probable end of support:
A | Handpiece:
1 | Should have speed 11000 CPM or more
2 | Should be light weight, not more than 2 kgs.
3 | Should be autoclavable.
4 | Should have safe mode.
5 | Should be provided with sternum saw attachment (If applicable) with blade guard.
B | Battery:
1 | Battery of NiMH type or Lithium Ion type
2 | Should have maximum charging time of 70 mins.
3 | Opening of battery housing should be easy for insertion of battery.
4 | Should have autoclavable or Non-autoclavable batteries- 02Nos.
5 | Should provide Battery casing incase of Non-autoclavable batteries-02 Nos.
C | Battery charger:
1 | should allow atleast 4 different batteries to be charged at the same time
2 | Condition of the battery should be easily assessed.
3 | charger should diagnose battery levels
D | Sterilization Kit
Should accommodate all handpieces, attachment and accessories for autoclave
List of consumables & accessories to be provided with quote.
International standards : The unit should comply with international standards and should have CE or FDA marking or BIS or ICMED.
Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABL norms). To submit SOPs for PMS at the time of installation/commissioning.
Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
Please specify pre installation requirements [electrical, HVAC, Compressed air ( please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour) ; if any of the above is essential]
Please specify footprint size & weight.
Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
Please specify power consumption.
Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
Warranty : 2 years
CMC : Maximum 5% per annum for 8 years on ex-works cost after warranty
Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted.
1. If 2 years warranty, CMC maximum 5% per annum for 8 years on ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3. Only Yes or No should be mentioned in compliance report and additional information should in remarks column.