## Sheet1
TMH/VIZ/2021-22/CAP/LT/0128
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Department - Surgery, HBCHRC Vizag.
Technical Specification of "Saline Warmer", Quantity 1 no.
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model:
Make:
Year of Introduction:
Probable end of support:
1 | System should be able to warm normal saline bottles of 500ml, 1000ml.
2 | Chamber should be of stainless steel with capacity of 75 litres or better.
3 | Door should be fully gasketed with a magnetic closure.
4 | Should have adjustable temperature from 30-65 degree celcius or better.
5 | Should have single door with glass windows with temperature reflective coating.
6 | Unit should be equipped with atleast two baskets to accommodate bottles or better.
7 | Should have capacity to keep minimum 20 nos. of 500ml bottles or better.
8 | Should have separate LED displays to display real time temperature and set temperature and both in Fahrenheit or Celcius.
9 | It should have accuracy +/- 1°C for set points.
10 | Should have an audible alarm if the temperature exceeds over the set point temperature.
11 | Should have an over temp indicator as well thermal shutdown as safety feature.
12 | Should be manufactured in an ISO certified facility and product should be FDA/CE/BIS/ICMED approved.
13 | Should work on 230V, 50/60Hz single phase electrical power
14 | Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABL norms). To submit SOPs for PMS at the time of installation/commissioning.
15 | Should be supplied with standard make UPS with atleast 30 mins backup (optional).
16 | Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
17 | Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
18 | Please specify pre installation requirements [electrical, HVAC, Compressed air ( please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour) ; if any of the above is essential]
19 | Please specify footprint size & weight.
20 | Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
21 | Pre-dispatch inspection at factory will be required if desired by authorities.
22 | Please visit the site to check whether the quoted model can be accomodated in the available space.
23 | Please specify power consumption.
24 | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
25 | Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
1 | Warranty : 2 years
2 | CMC : Maximum 5% per annum for 8 years on ex-works cost after warranty
3 | Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted.
1. If 2 years warranty, CMC maximum 5% per annum for 8 years on ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3. Only Yes or No should be mentioned in compliance report and additional information should in remarks column.