## Sheet1
TMH/VIZ/2021-22/CAP/LT/0129
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Department - HLA Lab, HBCHRC Vizag.
Technical Specification of "Deep Freezer -30 Degree C", Quantity 1 no.
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Year of Probable end of support:
1 | Capacity 400 to 500 litres.
2 | Temperature control range : -20°C to -30°C
3 | Temperature alarm : high / low ±5°C to ±15°C adjustable
4 | Power failure alarm : Buzzer and lamp, memory backup.
5 | Circular thermograph for continuous temperature monitoring for 7 days, inkless chart recorder.
6 | Should have green gas / CFC free refrigerant.
7 | Safety controls for system as well as compressor.
8 | Unit shall perform on full load at desired temperature levels.
9 | Locking arrangement to be provided.
10 | Alarm and SMS in case of temperature rise above acceptable limit shall be an added feature.
11 | Should be CE certified / BIS / FDA / ICMED approved.
12 | Power input 220V / 50Hz.
13 | Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards (As per NABL norms). To submit SOPs for PMS at the time of installation / commissioning.
14 | Performance certificates of IQ,OQ,PQ tests to be submitted at the time of installation/commissioning;if desired the end user (optional).
15 | Should be supplied with standard make UPS with atleast 30 mins backup (optional).
16 | Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
17 | Please specify list & cost of consumables/ consumable spares (i.e spares need to be replaced at regular intervals, may be quarterly/half yearly/ yearly such as annual maintenance kit etc.) if any.
18 | Please specify pre installation requirements (electrical, HVAC, Compressed air (Plaese specify required air pressure), water requirement (RO/ DI/ Distilled water with its pressuer, flow rate per hour); if any of the above is essential])
19 | Please specify footprint size & weight
20 | Demo of the quoted model, will be required, if so desired by the user, after the opening of technical bid and prior to the opening of financial bid. This is for technical evaluation.
21 | Predispatch inspection at factory will be required if desired by authorities.
22 | Please visit the site to check whether the quoted model can be accomodated in the available space.
23 | Please specify power consumption.
24 | Local Service Support : Should have local office and service support/service engineer for attending the breakdown calls.
26 | Response Time : Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
1 | Warranty : 2 years
2 | CMC : Maximum 5% per annum for 8 years on Ex-works cost after warranty.
3 | Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted.
1. If 2 years warranty, CMC maximum 5% per annum for 8 years on Ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3. Only Yes or No should be mentioned in compliance report and additional information in remarks column.