## Sheet1
TMH/TMH/2022-23/CAP/LT/0017
TATA MEMORIAL HOSPITAL
PURCHASE DEPARTMENT
Department - Transfusion Medicine, TMH.
Technical Specification of "Semi Automated Coagulometer", Quantity 1 no.
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Year of Probable end of support:
1 | Should be microcomputer controlled
2 | Semi automatic with at least 2-4 channels optics
3 | Based on opto-mechanical/viscosity based any other relevant measuring principle.
4 | Should have minimum 2 reagent,2 reading and 10 specimen positions at 37 ±10 C.
5 | Suitable for PT, a-PTT, fibrinogen, thrombin time, factors: II, V, VII, VIII, IX, X, XI, XII and other clot based assays.
6 | Results can be represented in seconds, % activity, ratio, INR g/L ang mg/L
7 | Should be able to store specific test parameters in the system and programmable assays.
8 | Should have LCD display
9 | Should provide Inbuilt roller/ticket printer/provision to attached to the computer and printer with software.
10 | Should generate the standard curve for factor assays
11 | System should be open so that any reagent can be used.
12 | Should have provision to feed sample ID.
13 | Should be able to perform QC, Levey Jennings plot and calibration of equipment.
14 | Should perform yearly calibration/Validation and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABL/NABH norms). To submit SOPs for PMS at the time of installation/commissioning.
15 | Performance certificates of IQ,OQ,PQ tests to be submitted at the time of installation/commissioning.
16 | Should be supplied with standard make UPS with atleast 30 mins backup.
17 | Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
18 | Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.(Quote Rate)
19 | Please specify pre installation requirements [electrical, HVAC, Compressed air ( please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour) ; if any of the above is essential]
20 | Please specify footprint size & weight.
21 | Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
22 | Predispatch inspection at factory will be required if desired by authorities.
23 | Please visit the site to check whether quoted model can be accomodated in the available space.
24 | Should be CE/FDA/BIS/ICMED approved
25 | Please specify power consumption.
26 | The Quoted model should be have minimum 10 installation in INDIA unless it is a new equipment - For eqpt more than 10 Lakhs
27 | Unpacking and shifting the consignment to the installation site is to be included in the scope of supply. Bidder/manufacturer/authorized service provider should take responsibility to lift /shift the consignment from unloading site to the installtion site. Unloading site shall be Stores department, Main Building Basement,TMH. If needed ,bidder has to arrange for the laboures at no charge to TMH.(Before submitting the quotation,bidders may visit TMH to Know unloading and installtion site.
28 | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
29 | Response Time: Should not be more than 06 hrs from lodging a breakdown complaint on toll free or by email.
30 | Warranty : 2 years
31 | AMC : Maximum 2% per annum for 8 years on Ex-works cost after warranty.
32 | Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted.
33 | Buyback Against old Model : CA 50, Make : Sysmex (Quantity 1 no.)
1. If 2 years warranty, AMC maximum 2% per annum for 8 years on Ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3. Only Yes or No should be mentioned in compliance report and additional information in remarks column.