## Sheet1
TMH/TMH/2022-23/CAP/LT/0023
TATA MEMORIAL HOSPITAL
PURCHASE DEPARTMENT
Department - Transfusion Medicine, TMH.
Technical specification of "Table Top Tube Sealer", Quantity 05 nos.
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No | Description | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Year of Probable end of support:
1 | Should seal PVC tubes of blood bags by radiofrequency (RF) for to avoid damage to blood inside the tubes without hemolysis
2 | Automated sensing system for sealing
3 | Easy separation of the sealed tube segment
4 | Protected sealing area/ protection against electric shock
5 | No warm up time and sealing time less than 2 seconds.
6 | Metal sealing head and separate rupture line to separate tube ends after sealing.
7 | Compatible with tubes of various manufacturers of blood bags.
8 | Should be able to seal 3-5 mm of tubes with wall thickness of 0.75 mm.
9 | Indication for READY, SEAL and POWER.
10 | Powercord length minimum 1.5 meter.
11 | Input 230-240 V AC 50/60 Hz
12 | Indication for cover open, power ready & seal prefered
13 | Automatic detection of tube by pressing of a lever which activates sensor
14 | Should perform yearly calibration/Validation and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABL/NABH norms). To submit SOPs for PMS at the time of installation/commissioning.
15 | Performance certificates of IQ,OQ,PQ tests to be submitted at the time of installation/commissioning.
16 | Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
17 | Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.(Quote Rate)
18 | Please specify pre installation requirements [electrical, HVAC, Compressed air ( please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour) ; if any of the above is essential]
19 | Please specify footprint size & weight.
20 | Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
21 | Predispatch inspection at factory will be required if desired by authorities.
22 | Please visit the site to check whether quoted model can be accomodated in the available space.
23 | Should be CE/FDA/BIS/ICMED approved
24 | Please specify power consumption.
25 | Unpacking and shifting the consignment to the installation site is to be included in the scope of supply. Bidder/manufacturer/authorized service provider should take responsibility to lift /shift the consignment from unloading site to the installtion site. Unloading site shall be Stores department, Main Building Basement,TMH. If needed ,bidder has to arrange for the laboures at no charge to TMH.(Before submitting the quotation,bidders may visist TMH to Know unloading and installtion site.
26 | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
27 | Response Time: Should not be more than 06 hrs from lodging a breakdown complaint on toll free or by email.
1 | Warranty : 2 years
2 | AMC : Maximum 2% per annum for 8 years on ex-works cost after warranty.
3 | Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted.
4 | Buyback offer: Against old hemopharm tube sealer- model BT 200 - 5 nos
1. If 2 years warranty, AMC maximum 2% per annum for 8 years on ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3. Only Yes or No should be mentioned in compliance report and additional information in remarks column.