## Sheet3
TMH/TMH/2022-23/CAP/LT/0024
TATA MEMORIAL HOSPITAL
PURCHASE DEPARTMENT
Department - Surgery , TMH.
Technical specification of "Portable OT Light", Quantity 03 nos.
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description | Complaince |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Year of Probable end of support:
1 | Should be a Mobile Surgical Light incorporating the latest LED technology only for Homogenous and shadow less operating light field.
2 | Should have mains cable and integrated ON/OFF switch on the light-head.
3 | It should have continuous height adjustment, weight balance in joint, no tilting, asymmetric stand , lockable castors.
4 | Should have Single Color LED’s with lifetime of 50000 hrs or more.
5 | The mobile light should have the following specifications :
a | Light Intensity 50000 lux or more
b | should have LEDs in the light head
c | Light Field diameter should be 20cm or better
d | Fixed Color temperature 4300 K or more
e | Color rendering index 95
f | Depth of Illumination 110 cm or better
6 | Should be provided with unique sterilisable (autoclavable) handle - 2 nos
7 | Should have stable illumination through out the life period of the light.
8 | Should perform yearly Preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABH norms) .To submit SOPs for PMS at the time of installation/commisioning.
9 | Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
10 | Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any. (Quote rate)
11 | Please specify pre installation requirements (electrical, HVAC etc.)
12 | Please specify footprint size & weight.
13 | Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
14 | Predispatch inspection at factory will be required if desired by authorities.
15 | Please visit the site to check whether quoted model can be accomodated in the available space.
16 | Should be CE/FDA/BIS/ICMED approved.
17 | Please specify power consumption.
18 | Unpacking and shifting the consignment to the installation site is to be included in the scope of supply. Bidder/manufacturer/authorized service provider should take responsibility to lift /shift the consignment from unloading site to the installtion site. Unloading site shall be Stores department, Main Building Basement,TMH. If needed ,bidder has to arrange for the laboures at no charge to TMH.(Before submitting the quotation,bidders may visist TMH to Know unloading and installtion site.
19 | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
20 | Response Time: Should not be more than 06 hrs from logging a breakdown complaint on toll free or by email.
1 | Warranty : 2 years
Warranty for the LEDs should be 10 years.
2 | AMC- Maximum 2% per annum for 8 years on ex-works cost after warranty.
Include 2 preventive maintenance/ year and all breakdown visits during warranty and AMC/CMC. All bought out items used in system should also be included in the warranty & AMC/CMC. Warranty & AMC/CMC should include upgrades / update cost, No separate software licence fees will be payable.
3 | Back to back assurance to be taken by the supplier from the OEM to supply spares for minimum 10 years and to be submitted.
4 | BUYBACK offer to be given for Existing Portable Light : 3Nos Make : Maquet Model : Blue 80
1. If 2 years warranty, AMC maximum 2% per annum for 8 years on ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3. Only Yes or No should be mentioned in compliance report and additional information in remarks column.