## Sheet1
TMH/TMH/2022-23/CAP/LT/0026
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Department - Medical Physics, TMH.
Technical specification of "Imported 1D Water Phantom", Quantity 2 nos.
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications.
Sr. No. | Parameters | Specifications | Compliance |  | Remarks
|  |  | Yes | No
Model:
Make:
Year of Introduction:
Year of Probable end of support:
1 | Application | Absolute Dose measurements
2 | Phanom wall Material & Density | Please specify (< 1.2 gm/cc)
3 | Size of Phantom | ≈ 42 cm (L) x 36 cm (W) x 36 cm (H)
| Inner dimensions | Please specify
| Exterior dimensions | Please specify
4 | Volume | Please specify
5 | Phantom leveling option | Please specify
6 | Sub milimeter scale on phantom | Please specify
7 | Manual movement of arm | Along Z-Axis (Vertical direction)
8 | Position indicator | Incremental encoder with Digital display
9 | Scanning mechanics | Noncorrosive
10 | Backscatter clearance at 25 cm depth | Please specify
11 | Scanning range | At least 25 cm
12 | Probe holder for detectors (existing detectors in the Department) | FC-65, CC13, NACP, Roos PP, Pin point 0.15cc, Semiflex, A1SL, Edge Diode detector
13 | Probe holder material | Water equivalent
14 | Positional Resolution | ± 0.1mm
15 | Positional Accuracy | ± 0.1mm
16 | Position Reproducibility | ± 0.1mm
17 | Power requirement | Please specify
18 | Water filling system | Please specify
19 | Water draining hose | Required, Please specify
20 | Weight | Please specify
21 | Virtual water slabs or Equivalent
| 30cm x 30cm x 0.1cm (Qty- 02) | Please specify
| 30cm x 30cm x 0.2cm (Qty- 04) | Please specify
| 30cm x 30cm x 0.5cm (Qty- 02) | Please specify
1 | Warranty | 2 years
1. If 2 years warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/ Stabilizer is required for the equipment. If so, should be included in the offer.
3. Only yes or no should be mentioned in the compilance report and additional info in remarks column.