## Sheet1
Serial # 4
TMH/TMH/2022-23/CAP/PT/0001
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Dept: Dept. Of Anaesthsia, Critical Care and Pain , TMH
Technical Specifications : Advanced Anesthesia Delivery System- Qty 05 nos
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Probable end of support:
1 | Durable with a sturdy kit.
2 | Antistatic wheels with brakes
3 | Shelves for monitors
4 | The anesthesia workstation should have in built ventilator with min. 12 inch TFT display.
5 | The anesthesia workstation should have electronic gas mixing siutable for adult, children & neanates.
6 | The anesthesia workstation should be suitable for low and minimal flow anesthesia application.
7 | The anesthesia workstation should have automatic calculations and presetting of patient specific ventilation settings via ideal body weight.
8 | It should have fully automatic menu drive self test and user check list. Incase of emergency it should be possible to bypass the self test.
9 | Gas delivery system:
a | Unit should have primary connection for central gas supply for Air, O2 & N2O and as a backup machine should have provision for connecting oxygen and nitrous pin index cylinders.
b | Machine should have elecronic fresh gas delivery & monitoring system with automatic switchover to air in case of failure of oxygen.
c | Audible and visual alarms for airway pressure, disconection, oxygen failure.
d | Anti-hypoxia safety mechanism to ensure that O2 concentration doesn’t fall below minimum 23% at any time.
e | O2 safety flow adjustable from 0 to 10L/min or more, runnig through the vaporiser for emergency backup use.
f | Fresh gas fow setting from 150ml /min(or less) to 12L/min
g | Electronic flowmeter for fresh gas flow.
h | Provision of additional O2 port for powering a ventury device.
i | Machine should also have an independent fresh gas outlet for connection to bain's or magill circuit.
10 | Breathing system:
a | Compact breathing system suitable for minimal flow anaesthesia with least patient circuit volume including absorber etc approx 3L (excluding bag) or less for fast response to change in fresh gas composition.
b | Fresh gas decoupled / compensated breathing system for adult & children.
c | should have mechanism to prevent moisture condunsation (eg. Heating circuit or equivalent )
11 | Ventilator
a | Ventilators for controlled ventilation of adults as well as children
b | Electronically controlled electrically or pneumatically driven ventilator.
c | Automatic breathing circuit compliance correction.
d | Provision for manual and mechanical ventilation
e | IPPV or equivalent with ability to adjust plateau time
f | pressure controlled ventilation.
g | SIMV with PCV & VCV
h | Pressure support, PS with CPAP, PS with PCV&VCV,SIMV with minimum Tidal volume upto 5ml.
i | High peak inspiratory flow upto 120LPM or more
j | Tidal volume delivery from 20 to 1400 ml or more
k | PEEP setting: upto 20 cmH20 or more
l | Respiration rate: 4(or less)-100 bpm
m | I:E : Max 2:1 to 1:8 or better
12 | Ventilator monitoring:
a | Monitoring of volume, pressure & oxygen
b | Should monitor patient compliance
c | should display curve & loops with fast analysis of change in lung mechanism provided by simultaneous display ETCO2,O2 as well as anaesthetic agent and automatic identification of agent .
d | Bar graph display of volumeter , tidal volume and the difference of inspiratory and expiratory tidal volume.
e | Trending for alarms and measured values with graphic trend display.
f | The machine should calculate agent cosumption on a case by case basis and display of fresh gas consuption in the unit logbook
13 | Alarms:
a | The machine should have adjustable alarm limits for all the parameters with auto set alarm function.
b | The machine should have automatic display of MAC values
c | Airway pressure low limit should automatically changes with the PEEP settings changes.
d | Machine should have RS232 connectivity port for interface to patient monitor/ HIS for automatic data acquisition.
14 | Vaporisers: (One Isoflurance and One Sevoflurance vaporiser with each machine)
a | temperature /pressure compensated and flow independent vaporizer.
b | Vaporizer should have extended delivery range from 0-6% volume
c | Vaporizer should have transport lock to provide hermetic sealing of agent chamber during transport & storage.
d | The vaporizer design should be maintenance free.
15 | Agent Monitoring:
a | should display agent analysis, N2O & MAC value with color coding of agents.
16 | Accessories
a | Should provide all standard acceossories
b | Open and closed circuit for both adult and pediatric
c | List of optional attachemts and accessories compatible with the specified model.
17 | Should work on mains Power and with battery backup of 45 minutes minimum with fully charged batteries
18 | Should be able to withstand temperature and humidity fluctuations in case of failure of air conditioning.
19 | Should perform yearly calibration/validation and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABL norms). To submit SOPs for PMS at the time of installation/commissioning.
20 | Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
21 | Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
22 | Please specify pre installation requirements [electrical, HVAC, Compressed air ( please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour) ; if any of the above is essential]
23 | Please specify footprint size & weight.
24 | Demo of the 'Quoted model' , will be mandatory, after the opening of the technical bid within ten days and prior to opening of financial bid. This is for technical evaluation.
25 | Predispatch inspection at factory will be required if desired by authorities.
26 | Please visit the site to check whether quoted model can be accomodated in the available space.
27 | Should be CE/FDA/BIS/ICMED approved
28 | Please specify power consumption.
29 | Unpacking and shifting the consignment to the installation site is to be included in the scope of supply. Bidder/manufacturer/authorized service provider should take responsibility to lift /shift the consignment from unloading site to the installtion site. Unloading site shall be Stores department, Main Building Basement,TMH. If needed ,bidder has to arrange for the laboures at no charge to TMH.(Before submitting the quotation,bidders may visist TMH to Know unloading and installtion site.
30 | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
31 | Response Time: Should not be more than 06 hrs from lodging a breakdown complaint on toll free or by email.
32 | Warranty- 2 years
33 | CMC- 5% of Ex-work cost per annum will be considered for all bidders for 8 years after warranty.
34 | Back to back assurance to be taken by the supplier from the OEM to supply spares for minimum 10 years and to be submitted to TMH .
Penalty Clause:
35 | The supplier and / or its Indian agent will be required to give the guarantee towards the performance of the equipment during the warranty period and in case of comprehensive annual maintenance contract period for maintaining the equipment in good working condition for a period of at least 300 days out of a period of 365 days a year (i.e. 95% uptime) 24 X 7. The ten hours non-functioning of the equipment or any part thereof will be considered as one day downtime. Total 3000 hours in a year will be considered as Uptime guarantee. Working time is 8.00a.m. To 6.00p.m. On week days and 8.00 a.m. To 4p.m. for Saturdays. The equipment shall be fully functional as part of the uptime and if it is partly functional it shall be considered as downtime and deduction shall be made on Pro-rata basis i.e. if the equipment is 70% functional & 30% downtime shall be applicable. The decision of the TMC management or its representative in determining the % of the downtime shall be final and binding. Essential period to shut down the installation entirely or partially shall also be included in the downtime while calculating the 95% guaranteed uptime i.e. all features as per specifications in purchase order should be functional for uptime. The supplier/agent shall be required to pay a penalty of 0.1% of total cost (Ex works) of equipment per day per machine excluding 150 hours permissible downtime. This rate shall increase commensurate with the rate increase of CMC. If the number of days downtime in each period of 365 days is more than the downtime permissible (which is to be calculated as defined above), the delay for not bringing the equipment in functioning order in any way, directly or even partially will be attributable to the firm. The bidder should provide the details of preventive maintenance to be undertaken through the year. Preventive maintenance to be done on holidays only. Preventive maintenance and breakdown will not be clubbed. The bill should be raised quarterly and will be paid at the end of every quarter only. If system is unserviceable for patients in its full capacity (full capacity : all features as demonstrated during commissioning) for more than 7 days at a stretch then these additional days exceeding 7 days will be considered as “downtime” and to levy penalty will be the discretion of Director TMC irrespective of 95% overall uptime in a year. This clause is to ensure maximum uninterrupted service to patients and hence Director, TMC’s decision in enforcing / invoking this clause will be final and binding for all. For CMC, W.O. issued by TMC shall be final as per D.A.E. Norms. TMC & its units shall not sign separate legal contract as per vendor’s format.
36 | Buyback Offer: Against old Narkomat Machine- 5 Nos
1) | If 2 years warranty and CMC 5% of ex works cost per annum will be considered for all bidders for 8 years after warranty . conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2) | Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3) | Only Yes or No should be mentioned in compliance report and additional information should in remarks column.