## Sheet1
Serial # 4
TMH/TMH/2022-23/CAP/PT/0002
TATA MEMORIAL HOSPITAL
PURCHASE DEPARTMENT
Dept- CSSD, TMH
Technical Specifications- ETO sterilizer Qty- 1
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sn | Item description | Compliance |  | Remarks
|  | Yes | No
Make:
Model:
Year of Introduction of the quoted model:
Year of Probable end of support:
1 | Should be fully microprocessor controlled sterilizer
2 | Should have 8 cubic ft chamber capacity with 100% ETO cartridge system
3 | Should operate at negative pressure in chamber compared to atmospheric pressure
4 | Should be compliant to all OSHA standard of USA and US FDA and CE marked or BIS or ICMED
5 | Should perform yearly calibration/Validation and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABH norms). To submit SOPs for PMS at the time of installation/commissioning.
6 | Performance certificates of IQ,OQ,PQ tests to be submitted at the time of installation/commissioning.
7 | Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
8 | Please specify list & cost of consumables/consumable Spares/annual maintenance kit i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly ) if any.(Quote Rate)
9 | Please specify pre installation requirements [electrical, HVAC, Compressed air ( please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour) ; if any of the above is essential]
10 | Please specify footprint size & weight.
11 | Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
12 | Predispatch inspection at factory will be required if desired by authorities.
13 | Please visit the site to check whether quoted model can be accomodated in the available space.
14 | Please specify power consumption.
15 | Unpacking and shifting the consignment to the installation site is to be included in the scope of supply. Bidder/manufacturer/authorized service provider should take responsibility to lift /shift the consignment from unloading site to the installtion site. Unloading site shall be Stores department, Main Building Basement,TMH. If needed ,bidder has to arrange for the laboures at no charge to TMH.(Before submitting the quotation,bidders may visist TMH to Know unloading and installtion site.
16 | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
17 | Response Time: Should not be more than 06 hrs from lodging a breakdown complaint on toll free or by email.
18 | Quoted company should be ISO certified.
19 | PENALTY CLAUSE:-
The supplier and / or its Indian agent will be required to give the guarantee towards the performance of the equipment during the warranty period and in case of comprehensive annual maintenance contract period for maintaining the equipment in good working condition for a period of at least 300 days out of a period of 365 days a year (i.e. 95% uptime) 24 X 7. The ten hours non-functioning of the equipment or any part thereof will be considered as one day downtime. Total 3000 hours in a year will be considered as Uptime guarantee. Working time is 8.00a.m. To 6.00p.m. On week days and 8.00 a.m. To 4p.m. for Saturdays. The equipment shall be fully functional as part of the uptime and if it is partly functional it shall be considered as downtime and deduction shall be made on Pro-rata basis i.e. if the equipment is 70% functional & 30% downtime shall be applicable. The decision of the TMC management or its representative in determining the % of the downtime shall be final and binding. Essential period to shut down the installation entirely or partially shall also be included in the downtime while calculating the 95% guaranteed uptime i.e. all features as per specifications in purchase order should be functional for uptime. The supplier/agent shall be required to pay a penalty of 0.1% of total cost (Ex works) of equipment per day per machine excluding 150 hours permissible downtime. This rate shall increase commensurate with the rate increase of CMC. If the number of days downtime in each period of 365 days is more than the downtime permissible (which is to be calculated as defined above), the delay for not bringing the equipment in functioning order in any way, directly or even partially will be attributable to the firm. The bidder should provide the details of preventive maintenance to be undertaken through the year. Preventive maintenance to be done on holidays only. Preventive maintenance and breakdown will not be clubbed. The bill should be raised quarterly and will be paid at the end of every quarter only.
1 | Warranty- 2 years
2 | CMC- 5% of Ex-works cost per annum will be considered for all bidders for 8 years after warranty. CMC should include cost for upgrades and updates cost, No separate software licence fees will be payable.
3 | Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted.
4 | Buy Back Offer- Against 1 ETO sterilizer of 3M make
1) | If 2 years warranty and CMC maximum 5% of Ex works cost per annum will be consided for 8 years after warranty . conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2) | Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3) | Only Yes or No should be mentioned in compliance report and additional information should in remarks column.