## Sheet1
Serial # 4
TMH/TMH/2022-23/CAP/PT/0003
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Department - Nursing and Anaesthesia, TMH.
Technical Specification of "Infusion Pumps", Quantity 51 nos. (Qty: 30 Nos for Nursing Dept. And Qty: 21 Nos for Aneasthesia)
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Year of Probable end of support:
1 | Delivery of medications indicated for infusion therapy, mains and battery operation.
2 | Comprehensive alarm system for following parameters or similar should be available:
~ Occlusion,
~ end of infusion,
~ lock mode,
~ air in line,
~ door open,
~ low battery,
~ system malfunction, etc
3 | Intended for use on adult/pediatric.
4 | Easy dose rate calculation, easy loading and operation process.
5 | Should have drug library.
6 | Should have volume and drop control.
7 | Equipment should fit on IV stand, should be compact.
8 | Universal plugging system and should fit in the present electric board.
9 | Equipment should be compatible with all standard IV sets including paediatric IV sets
10 | Battery charging time should be less than 8 hours for depleted battery.
11 | Battery back up should be up to 4 hrs at 25 ml/hr
12 | Pole clamp material should be of metal or hard teflon.
13 | Should have drip sensor.
14 | Should have purge function.
15 | Infusion Volume limit should be upto to 999.9ml
16 | Over all accuracy of flow should be ± 10% or better.
17 | Should have bright LED/LCD display showing infusion parameters.
18 | Air bubble sensor should be available
19 | Equipment should be capable for preventing free flow and back flow.
20 | Should be light weight, pls mention weight.
21 | Should perform yearly calibration/Validation and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABL/NABH norms). To submit SOPs for PMS at the time of installation/commissioning.
22 | Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
23 | Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.(Quote Rate)
24 | Please specify pre installation requirements [electrical, HVAC, Compressed air ( please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour) ; if any of the above is essential]
25 | Please specify footprint size & weight.
26 | Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
27 | Predispatch inspection at factory will be required if desired by authorities.
28 | Please visit the site to check whether quoted model can be accomodated in the available space.
29 | Should be CE/FDA/BIS/ICMED approved
30 | Please specify power consumption.
31 | Unpacking and shifting the consignment to the installation site is to be included in the scope of supply. Bidder/manufacturer/authorized service provider should take responsibility to lift /shift the consignment from unloading site to the installtion site. Unloading site shall be Stores department, Main Building Basement,TMH. If needed ,bidder has to arrange for the laboures at no charge to TMH.(Before submitting the quotation,bidders may visist TMH to Know unloading and installtion site.
32 | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
33 | Response Time: Should not be more than 06 hrs from lodging a breakdown complaint on toll free or by email.
1 | Warranty : 2 years
2 | AMC : Maximum 2% per annum for 8 years on ex-works cost after warranty
3 | Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted.
4 | Buyback against old Model : 1. Infusomat Space, Make : B Braun (Qty. 14 nos.) 2. Model : OT 701, Make : JMS Pumps (Qty. 7 nos.) 3. Make B Braun Qty: 30 Nos
1. If 2 years warranty, AMC maximum 2% per annum for 8 years on ex-works cost after warranty and buyback conditions wherever applicable are not adhered to TMC, reserves right to reject the offer.
2. Please mention whether UPS/ stabiliser is required for the equipment. If so should be included in the offer.
3. Only Yes or No should be mentioned in compliance report and additional information in remarks column.