## Sheet1
Serial # 4
TMH/TMH/2022-23/CAP/PT/0004
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Dept: Dept. Of Anaesthsia, Critical Care and Pain , TMH
Technical Specifications : Anesthesia Delivery System - Basic (Qty: 07 nos)
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Complieance |  | Remark
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Probable end of support:
1 | Durable with a sturdy kit.
2 | Antistatic wheels with brakes
3 | Shelves for monitors
4 | The anesthesia workstation should have in built ventilator with min. 6'' TFT display.
5 | Gas delivery system:
a | Gas specific high pressure blocks.
b | Non interchangable gas supply inlet for cylinders with step down regulators
c | Pressure reducing valves.
d | Color coded indexed gas supply hoses.
e | Provision of additional O2 port for powering a ventury device.(Optional)
f | Provision for manual as well as mechanical ventilation.
g | Provision for both open as well as closed circuit.
h | Provision for attachement of 2 cylinders one each of O2 and N2O cylinders.
i | Provision to monitor the expired gas volume, O2 monitor sensor probe along with all the tubings and connections with alarms.
j | Ability to monitor the expired TV/minute volume,respiratoy rate and FiO2
k | Quick change over form open circuit to closed cicuit and vice versa.
l | Ability to bypass the soda lime canister and convert to an open circuit.
m | Provision for emergency flow of 100% O2 under high pressure through both circuits.
n | Audible and visual alarms for airway pressure, disconection, oxygen failure.
6 | Flow meters:
a | For oxygen, nitrous oxide and air with safety feature
b | Oxygen failure alarm, audible and visible
c | Anti-hypoxia safety mechanism to ensure that O2 concentration doesn’t fall below minimum 23% at any time.
7 | Vaporisers (One Isoflurane and One sevoflurane vaporiser with each machine)
a | Should have provision to connect two selectatec mount vaporisers.
b | Vaporiser safety lock.
8 | Ventilators:
a | Ventilators for controlled ventilation of adults as well as children
b | Volume contorlled/pressure controlled/Pressure support
c | Inbuilt PEEP administration.
d | Provision for manual and mechanical ventilation
e | Monitoring for airway prsesure- Peak and PEEP
f | Provision to monitor expired gas volume, O2 monitor sensor probe along with all the tubings and connections with alarms.
g | Ability to monitor the expired TV/minute volume, tidal volume, respiratoy rate and FiO2
h | Tidal volume delivery should be in the range from 30ml(or less) to 1400 ml (or more)
i | Respiration rate: 4-60 bpm (or more)
j | PEEP setting: upto 20 cmH20
k | Peak pressure setting upto 50cmH20
9 | Breathing system:
a | Should have integrated CO2 absorber with easy removal and refitting during the operation.
b | Should be autoclavable at 134 deg C.
c | Should have APL valve, inspiratory valve, expiratory valve.
d | should have mechanism to prevent moisture condensation (eg. Heating circuit or equivalent )
10 | Should work on mains Power and with battery backup of 45 minutes minimum with fully charged batteries
11 | Accessories
a | Should provide all standard acceossories
b | Open and closed circuit for both adult and pediatric
c | List of optional attachments and accessories compatible with the specified model.
d | Should provide additional soda lime canisters - 2 nos
Should perform yearly calibration/Validation and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABL/NABH norms). To submit SOPs for PMS at the time of installation/commissioning.
Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.(Quote Rate)
Please specify pre installation requirements [electrical, HVAC, Compressed air ( please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour) ; if any of the above is essential]
Please specify footprint size & weight.
Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
Predispatch inspection at factory will be required if desired by authorities.
Please visit the site to check whether quoted model can be accomodated in the available space.
Should be CE/FDA/BIS/ICMED approved
Please specify power consumption.
Unpacking and shifting the consignment to the installation site is to be included in the scope of supply. Bidder/manufacturer/authorized service provider should take responsibility to lift /shift the consignment from unloading site to the installtion site. Unloading site shall be Stores department, Main Building Basement,TMH. If needed ,bidder has to arrange for the laboures at no charge to TMH.(Before submitting the quotation,bidders may visist TMH to Know unloading and installtion site.
Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
Response Time: Should not be more than 06 hrs from lodging a breakdown complaint on toll free or by email.
Warranty- 2 years
CMC- 5% of Ex-works cost per annum will be considered for all bidders for 8 years after warranty
Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted.
Buyback offer: Against old Narkomat machines - 7 nos
1) | If 2 years warranty and CMC 5% of ex works cost per annum will be considered for all bidders for 8 years after warranty . conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2) | Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3) | Only Yes or No should be mentioned in compliance report and additional information should in remarks column.