## Tech Specs
Serial #04
TMH/TMC/2022-23/CAP/PT/0006
TATA MEMORIAL HOSPITAL
PURCHASE DEPARTMENT
Dept Name:Surgery, HBCHRC, Vizag and HBCHRC Mullanpur
Technical specifications - "Micromotor drill with saw", QTY: 05 Nos (Qty': 02 Nos for HBCHRC Vizag and qty: 03 Nos for HBCHRC Mulanpur)
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Year of Probable end of support:
1 | Electric Console
a | Should have integrated console drill and saw system.
b | Should be able to identify different handpieces with display on control.
c | Torque sensing feedback capability.
d | Console should offer change from variable/fixed mode, forward/reverese/osscilation mode.
e | Console should offer change of speed.
f | Console should offer change in irrigation percentage.
g | Integrated irrigation pump should be available for drill & saw system.
2 | Footswitch - 1no
It should have:
a | Should have fully programmable footswitch as user needs
b | User should be able to control following functions via footswitch:
> Forward,reverse, oscillation
3 | High speed micro drill (with straight & angled attachment)
It should have:
a | Maintenance free, brushless motor
b | Handpiece speed should be atleast 70,000 RPM or more
c | Sterilizable through steam, ETO, flash autoclavable.
4 | Attachments required:
a | 7.5 CMS STRAIGHT ATTACHMENT or equivalent
b | 7.5 CMS ANGLED ATTACHMENT or equivalent
c | 15.0 CMS ANGLED ATTACHMENT or equivalent
5 | Micro Saw :
a | maintenance free D.C brushless motor, to be provided independently with each type of saw.
b | Should be autoclavable
c | locking mechanism in order to avoid any mishaps
d | Should be provided with sufficient nos of compatible blades
e | Oscillating saw
a. Should cut at right angle
b. Speed should be 16000cpm or more
f | Sagittal Saw
a. Should cut in horizontal plane
b. Speed should be 20000cpm or more
g | Reciprocating Saw :
a. speed should be 14000cpm or more
6 | Connecting Cord
1 | Should be 10ft long, 3/8'' diameter flexible electrical connecting cord.
2 | Should be autoclavable
Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABH norms)
Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
Please specify pre installation requirements (electrical, HVAC etc.)
Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
Predispatch inspection at factory will be required if desired by authorities.
Please specify power consumption.
The unit should comply with CE/ ICMED/BIS
Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
Warranty: 2 years
CMC: CMC @maximum 5% per annum of ex-works cost of the equipment will be considered for all bidders for 8 years after warranty
Back to back assurance to be taken by the supplier from the OEM to supply spares for minimum 10 years and to be submitted
1) If 2 years warranty, CMC @maximum 5% per annum of ex-works cost of the equipment will be considered for all bidders for 8 years after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject offer.
2) Please mention whether UPS / Stabilizer is required for the equipment. If so, should be included in the offer.
3) Only yes or no should be mentioned in the compliance report and additional info in remarks column.