## ICU
Serial # 4
TMH/TMC/2022-23/CAP/PT/0008
TATA MEMORIAL HOSPITAL
PURCHASE DEPARTMENT
DEPARTMENT: ICU, HBCHRC Vizag and HBCHRC Mullanpur
Technical Specifications of "Portable Ultasound with cardiac & Linear Probe " OTY : 02 Nos (Qty: 01 no each for Vizag and Mullanpur)
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item |  | Compliance |  | Remarks
| Specifications | Detailed Specifications | Yes | No
Model Name :
Make :
Year of Introduction :
Probable end of support :
1 | Beam Former | Full Digital Beam Former
2 | Imaging Modes | a) 2D, Dual 2D imaging
b) M-Mode with variable sweep speed
c) Steerable PW Doppler with angle corrections
d) Color Flow Mapping with various maps
e) Power Doppler with at least 3 different maps.
f) Chroma color or B-color with various maps
3 | Composite Modes | B/M, B/PW, B/CW, B/CFM/M, B/CFM/PW (Triplex Mode)
4 | Image depth | System should have depth upto 30cm or more
5 | Transducer Characteristics | a) Probe range should be 2 MHz to 13 MHz frequency Range. (Probe dependent). Specify the frequency range.
b) Each probe should have three operating frequencies.
6 | Applications | a) General Abdomen & Urology
b) Adult Cardiac
c) Pediatric Cardiac
b) Musculoskeletal
c) Vascular/ Neural ( for nerve block)
d) Small parts (breast, thyroid, ophthalmology, Scrotum)
e) Pediatrics
7 | Monitor Type | Digital High Resolution
8 | Monitor Size | Minimum 15 Inch or more
9 | Dynamic Range | System should have minimum dynamic range approx upto 200db.
10 | Probe Connectors | a) Minimum 3 inuilt active connectors for adult cardiac , peadiatric cardiac, TEE, Convex, Linear, TV+TR, Intra-operative probes or more
b) Probe & Frequency selection should be through electronic switch available on the console of the unit.
11 | Battery and UPS | Please Provide Emerson make online UPS with isolation transformer of adequte capacity for the complete systems. With 30 min back up.
12 | Weight | System should weight maximum 8kg
13 | Cine Loop Frames | More than 300 loops review (variable duration)
14 | Memory Storage | The system should be capable to have memroy storage of 250 GB internal or external.
15 | Zoom | The system should have Zoom on real time & frozen images.
16 | General Measurement & calculations for B, M & Doppler Modes | a) B-Mode: Distance, ratio distance, stenosis distance, length by trace, area by ellipse axis, area by vertex, area by trace, ratio area, stenosis area, volume by ellipse axis, volume by trace.
b) M-Mode: distance, ratio distance, time, ratio time, heart rate, velocity, ratio velocity, slope.
c) PW-Mode: Trace, time, ratio time, velocity, ratio velocity, heart rate, S/D, RI, slope.real time Doppler measurements.
17 | Tissue Harmonic Imaging- THI | Must be available on all probes to improve contrast resolution & reduces Image artifacts in obese or difficult patients.
18 | Speckle Reduction Imaging | Machine must have facility of speckle reduction (Noise & artifacts) imaging technique as standard with 2-3 modes for best visualization
19 | Special compound Imaging technique with linear probes | Machine must have facility of compound / cross beam technique to reduce artifacts, shadow and speeckle for best visualization of Linear probes imaging.
20 | Doppler Functions | User selectable parameters for PRF, Wall Filters, Sample Size, Velocity, Frequency, angle Corrections, Dynamic Range & post processing.
21 | Real Time Doppler Measurements | RI & PI real time measurement must be possible to be measured online.
22 | Doppler Audio | Audio output with speakers
23 | Color Flow Mapping | User selectable parameters for PRF,Color filters, Velocity (Mean, Velocity / variance, Variance & Intensity etc. Maps) Variable CFM Box, Color Processing.
24 | Tissue doppler imaging | Should be provided
25 |  | Should have upgradibility to perform advanced quantification measurements like strain and strain rate quantification.
26 | Steering of Box | 20 degree in a Linear probe Application.
27 | Simultaneous Mode | B + (B with CFM in one image), both must be lvie at a time for better evaluation.
28 | Archiving facility
| Hard Disc Drive | a) Machine must have facility to transfer & store Images, Loops & data on Hard disc drive
| DVD writer | b)Machine must have facility to transfer/ write Images, Loops & Data on internal or external DVD writer
| USB | c) Machine must have facility of at least 2 Usb ports to transfer/store Images, Loops on USB pen drive
| Compatibility with our existing PACS | d) Compatibility with our existing PaCS, Dicom send/receive, query/retrieve, basic print, HIS/RIS work list, storage, MPPS, Storage commit.
e) System should have DICOM print and storage class software.
29 | Measurement & calculation any Computer | Measurements must be possible on Frozen Images and cine loops which are stored on CD/DVD & USB pen drive.
30 | Very High frequency imaging facility | Machine must have capability to support very high frequency imaging 15 MHz or more.
31 | Trolley | The machine should be supplied with trolley of same make.
32 | Foot Switch | The machine should be supplied with foot switch
33 | Thermal Printer | The machine should be supplied with thermal printer
34 | System should have FDA/ CE/ BIS/ICMED approval.
35 | The protective membrane or protective lens of the probe should cover only top side of the probe not the side portion.
36 | Probes to be supplied with Machine
i) | Cardiac Probe | 1 to 4 Mhz Broad Band Probe
ii) | Linear Probe | 4 to 13 MHz Broad Band Probe
iii) | Convex Probe | 2 to 5 MHz Broad Band probe, Single Crystal/ XD Clear/Purewave Convex Probe.
iv) | System should be supplied with ECG module & its accessories
37 | Performance certificates of IQ,OQ,PQ tests to be submitted at the time of installation/commissioning ; if desired by the end user. (Optional)
38 | Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
39 | Please specify pre installation requirements [electrical, HVAC, Compressed air ( please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour) ; if any of the above is essential]
40 | Please specify footprint size & weight.
41 | Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
42 | Predispatch inspection at factory will be required if desired by authorities.
43 | Please visit the site to check whether quoted model can be accomodated in the available space.
44 | Please specify power consumption.
45 | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
46 | Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
47 | Warranty : 2 Years
48 | CMC :CMC maximum 5% of Ex works cost per annum will be consided for 8 years after warranty
49 | CMC to be included with probes.
50 | PENALTY CLAUSE
The supplier and / or its Indian agent will be required to give the guarantee towards the performance of the equipment during the warranty period and in case of comprehensive annual maintenance contract period for maintaining the equipment in good working condition for a period of at least 300 days out of a period of 365 days a year (i.e. 95% uptime) 24 X 7. The ten hours non-functioning of the equipment or any part thereof will be considered as one day downtime. Total 3000 hours in a year will be considered as Uptime guarantee. Working time is 8.00a.m. To 6.00p.m. On week days and 8.00 a.m. To 4p.m. for Saturdays. The equipment shall be fully functional as part of the uptime and if it is partly functional it shall be considered as downtime and deduction shall be made on Pro-rata basis i.e. if the equipment is 70% functional & 30% downtime shall be applicable. The decision of the TMC management or its representative in determining the % of the downtime shall be final and binding. Essential period to shut down the installation entirely or partially shall also be included in the downtime while calculating the 95% guaranteed uptime i.e. all features as per specifications in purchase order should be functional for uptime. The supplier/agent shall be required to pay a penalty of 0.1% of total cost (Ex works) of equipment per day per machine excluding 150 hours permissible downtime. This rate shall increase commensurate with the rate increase of CMC. If the number of days downtime in each period of 365 days is more than the downtime permissible (which is to be calculated as defined above), the delay for not bringing the equipment in functioning order in any way, directly or even partially will be attributable to the firm. The bidder should provide the details of preventive maintenance to be undertaken through the year. Preventive maintenance to be done on holidays only. Preventive maintenance and breakdown will not be clubbed. The bill should be raised quarterly and will be paid at the end of every quarter only. If system is unserviceable for patients in its full capacity (full capacity : all features as demonstrated during commissioning) for more than 7 days at a stretch then these additional days exceeding 7 days will be considered as “downtime” and to levy penalty will be the discretion of Director TMC irrespective of 95% overall uptime in a year. This clause is to ensure maximum uninterrupted service to patients and hence Director, TMC’s decision in enforcing / invoking this clause will be final and binding for all. For CMC, W.O. issued by TMC shall be final as per D.A.E. Norms. TMC & its units shall not sign separate legal contract as per vendor’s format.
51 | Back to back assurance to be taken by the supplier from the OEM to supply spares for minimum 10 years and to be submitted.
1) | If 2 years warranty and CMC maximum 5% of Ex works cost per annum will be consided for 8 years after warranty . conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2) | Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3) | Only Yes or No should be mentioned in compliance report and additional information should in remarks column.
Note: Please refer 2nd worksheet for technical specification of Radiology Dept, HBCHRC Mullanpur

## Radiology
Serial # 4
TMH/TMC/2022-23/CAP/PT/0008
TATA MEMORIAL HOSPITAL
PURCHASE DEPARTMENT
DEPARTMENT: Radiology, HBCHRC Mullanpur
Technical Specifications of "Portable Ultasound with cardiac & Linear Probe " OTY : 04 Nos
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item |  | Compliance |  | Remarks
| Specifications | Detailed Specifications | Yes | No
Model Name :
Make :
Year of Introduction :
Probable end of support :
1 | Beam Former | Full Digital Beam Former
2 | Imaging Modes | a) 2D, Dual 2D imaging
b) M-Mode with variable sweep speed
c) Steerable PW Doppler with angle corrections
d) Color Flow Mapping with various maps
e) Power Doppler with at least 3 different maps.
f) Chroma color or B-color with various maps
3 | Composite Modes | B/M, B/PW, B/CW, B/CFM/M, B/CFM/PW (Triplex Mode)
4 | Image depth | System should have depth upto 30cm or more
5 | Transducer Characteristics | a) Probe range should be 2 MHz to 13 MHz frequency Range. (Probe dependent). Specify the frequency range.
b) Each probe should have three operating frequencies.
6 | Applications | a) General Abdomen & Urology
b) Musculoskeletal
c) Vascular
d) Small parts (breast, thyroid, ophthalmology, Scrotum)
e) Pediatrics
7 | Monitor Type | Digital High Resolution
8 | Monitor Size | Minimum 15 Inch or more
9 | Dynamic Range | System should have minimum dynamic range approx upto 200db.
10 | Probe Connectors | a) Minimum 3 inuilt active connectors for adult cardiac , peadiatric cardiac, TEE, Convex, Linear, TV+TR, Intra-operative probes or more
b) Probe & Frequency selection should be through electronic switch available on the console of the unit.
11 | Battery and UPS | Please Provide online UPS with isolation transformer of adequte capacity for the complete systems. With 30 min back up.
12 | Weight | System should weight maximum 8kg
13 | Cine Loop Frames | More than 300 loops review (variable duration)
14 | Memory Storage | The system should be capable to have memroy storage of 250 GB internal or external.
15 | Zoom | The system should have Zoom on real time & frozen images.
16 | Trapezoidal Lamaging | Linear Probe must have facility of Trapezoidal Imaging.
17 | Panoramic Imaging | Machine must have capability of Panoramic Imaging
18 | General Measurement & calculations for B, M & Doppler Modes | a) B-Mode: Distance, ratio distance, stenosis distance, length by trace, area by ellipse axis, area by vertex, area by trace, ratio area, stenosis area, volume by ellipse axis, volume by trace.
b) M-Mode: distance, ratio distance, time, ratio time, heart rate, velocity, ratio velocity, slope.
c) PW-Mode: Trace, time, ratio time, velocity, ratio velocity, heart rate, S/D, RI, slope.real time Doppler measurements.
19 | Tissue Harmonic Imaging- THI | Must be available on all probes to improve contrast resolution & reduces Image artifacts in obese or difficult patients.
20 | Speckle Reduction Imaging | Machine must have facility of speckle reduction (Noise & artifacts) imaging technique as standard with 2-3 modes for best visualization
21 | Special compound Imaging technique with linear probes | Machine must have facility of compound / cross beam technique to reduce artifacts, shadow and speeckle for best visualization of Linear probes imaging.
22 | Doppler Functions | User selectable parameters for PRF, Wall Filters, Sample Size, Velocity, Frequency, angle Corrections, Dynamic Range & post processing.
23 | Real Time Doppler Measurements | RI & PI real time measurement must be possible to be measured online.
24 | Doppler Audio | Audio output with speakers
25 | Color Flow Mapping | User selectable parameters for PRF,Color filters, Velocity (Mean, Velocity / variance, Variance & Intensity etc. Maps) Variable CFM Box, Color Processing.
26 | Tissue doppler imaging | Should be provided
27 | Steering of Box | 20 degree in a Linear probe Application.
28 | Simultaneous Mode | B + (B with CFM in one image), both must be lvie at a time for better evaluation.
29 | Archiving facility
| Hard Disc Drive | a) Machine must have facility to transfer & store Images, Loops & data on Hard disc drive
| DVD writer | b)Machine must have facility to transfer/ write Images, Loops & Data on internal DVD or ecternal DVD writer
| USB | c) Machine must have facility of at least 2 Usb ports to transfer/store Images, Loops on USB pen drive
| Compatibility with our existing PACS | d) Compatibility with our existing PaCS, Dicom send/receive, query/retrieve, basic print, HIS/RIS work list, storage, MPPS, Storage commit.
e) System should have DICOM print and storage class software.
30 | Measurement & calculation any Computer | Measurements must be possible on Frozen Images and cine loops which are stored on CD/DVD & USB pen drive.
31 | Very High frequency imaging facility | Machine must have capability to support very high frequency imaging 15 MHz or more.
32 | Trolley | The machine should be supplied with trolley of same make.
33 | Foot Switch | The machine should be supplied with foot switch
34 | Thermal Printer | The machine should be supplied with thermal printer
35 | System should have US FDA/ CE/ BIS/ICMED approval.
36 | TV/TR Probe with biopsy attachement kit/package
37 | The protective membrane or protective lens of the probe should cover only top side of the probe not the side portion.
38 | Probes to be supplied with Machine
i) | TV/TR Probe | 4 to 10 Mhz Broad Band Probe
ii) | Linear Probe | 4 to 13 MHz Broad Band Probe
iii) | Convex Probe | 2 to 5 MHz Broad Band probe, Single Crystal/ XD Clear/Purewave Convex Probe.
iv) | System should be supplied with ECG module & its accessories
39 | Performance certificates of IQ,OQ,PQ tests to be submitted at the time of installation/commissioning ; if desired by the end user. (Optional)
40 | Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
41 | Please specify pre installation requirements [electrical, HVAC, Compressed air ( please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour) ; if any of the above is essential]
42 | Please specify footprint size & weight.
43 | Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
44 | Predispatch inspection at factory will be required if desired by authorities.
45 | Please visit the site to check whether quoted model can be accomodated in the available space.
46 | Please specify power consumption.
47 | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
48 | Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
49 | Warranty : 2 Years
CMC :CMC maximum 5% of Ex works cost per annum will be consided for 8 years after warranty
CMC to be included with probes.
50 | PENALTY CLAUSE
The supplier and / or its Indian agent will be required to give the guarantee towards the performance of the equipment during the warranty period and in case of comprehensive annual maintenance contract period for maintaining the equipment in good working condition for a period of at least 300 days out of a period of 365 days a year (i.e. 95% uptime) 24 X 7. The ten hours non-functioning of the equipment or any part thereof will be considered as one day downtime. Total 3000 hours in a year will be considered as Uptime guarantee. Working time is 8.00a.m. To 6.00p.m. On week days and 8.00 a.m. To 4p.m. for Saturdays. The equipment shall be fully functional as part of the uptime and if it is partly functional it shall be considered as downtime and deduction shall be made on Pro-rata basis i.e. if the equipment is 70% functional & 30% downtime shall be applicable. The decision of the TMC management or its representative in determining the % of the downtime shall be final and binding. Essential period to shut down the installation entirely or partially shall also be included in the downtime while calculating the 95% guaranteed uptime i.e. all features as per specifications in purchase order should be functional for uptime. The supplier/agent shall be required to pay a penalty of 0.1% of total cost (Ex works) of equipment per day per machine excluding 150 hours permissible downtime. This rate shall increase commensurate with the rate increase of CMC. If the number of days downtime in each period of 365 days is more than the downtime permissible (which is to be calculated as defined above), the delay for not bringing the equipment in functioning order in any way, directly or even partially will be attributable to the firm. The bidder should provide the details of preventive maintenance to be undertaken through the year. Preventive maintenance to be done on holidays only. Preventive maintenance and breakdown will not be clubbed. The bill should be raised quarterly and will be paid at the end of every quarter only. If system is unserviceable for patients in its full capacity (full capacity : all features as demonstrated during commissioning) for more than 7 days at a stretch then these additional days exceeding 7 days will be considered as “downtime” and to levy penalty will be the discretion of Director TMC irrespective of 95% overall uptime in a year. This clause is to ensure maximum uninterrupted service to patients and hence Director, TMC’s decision in enforcing / invoking this clause will be final and binding for all. For CMC, W.O. issued by TMC shall be final as per D.A.E. Norms. TMC & its units shall not sign separate legal contract as per vendor’s format.
51 | Back to back assurance to be taken by the supplier from the OEM to supply spares for minimum 10 years and to be submitted.
1) | If 2 years warranty and CMC maximum 5% of Ex works cost per annum will be consided for 8 years after warranty . conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2) | Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3) | Only Yes or No should be mentioned in compliance report and additional information should in remarks column.