## Sheet1
Serial # 4
TMH/MUL/2022-23/CAP/PT/0009
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Dept Name: General Medicine, HBCHRC Mullanpur
Technical Specifications: ECG machine - 12 nos
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Year of Probable end of support:
1 | Should be Fully automatic, microprocessor controlled mains and battery operated.
2 | All 12 leads of ECG can be measured and analyzed simultaneously.
3 | Should have either backlit LCD or color TFT display.
4 | Should have detailed measurement results table, pacemaker detection.
5 | Should have inbuilt high resolution printer.
6 | Should be portable and light weight.
7 | Should have printer paper with A4 size as standard.
8 | ECG Interpretation should be available
9 | Alpha Numeric keyboard facility for Patient data entry –Name, Age, Patient ID, Sex, Weight should be available
10 | All type of noise filter option should be available.
11 | Rechargeable battery operation with 100 ECG recordings on fully charged battery or 3 hrs of continuous operation.
12 | Battery charging and Low battery indication should be available.
13 | 25, 50mm/s Paper speed should be available.
14 | Should have Indication for bad electrode placement or similar.
15 | Should have Memory for ECG storage should be atleast 40 ECGs.
16 | Should have FDA/CE/BIS certifications.
17 | Should have facility to transfer the ECG reports to any 'N' no of desktop computers without any additional software requirement which then can be displayed on EMR in PDF. This facility will have to be demonstated at the time of equipment demo.
18 | It should have capability to transfer the ECG data in XML, DICOM format.
19 | Should be supplied with Interfacing software for data transmission and It should have DICOM compatibility
20 | Standard accessories :
a | 10 lead ECG cable -1 no
b | Standard Limb electrodes -1 no
c | Chest electrodes with suitable connectors
d | ECG electrodes limb electrodes and chest electrodes
e | ECG paper – 10 packs
21 | Should perform yearly calibration/validation and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards (As per NABH norms). To submit SOPs for PMS at the time of installation/commissioning.
22 | Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
23 | Please specify list & cost of consumables/consumable Spares (i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
24 | Please specify pre installation requirements [electrical, HVAC, Compressed air (please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour) ; if any of the above is essential]
25 | Please specify footprint size & weight.
26 | Demo of the 'Quoted model' , will be mandatory, after the opening of the technical bid within ten days and prior to opening of financial bid. This is for technical evaluation.
27 | Please specify power consumption.
28 | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls. Give details of office and contact nos of service personnel locally available
29 | Response Time: Should not be more than 06 hrs from lodging a breakdown complaint on toll free or by email.
Warranty- 2 years (If not responded to breakdown call as per the response time of 6 hrs vendor will have to pay penalty of Rs. 1000/-, per day during warranty)
AMC- 2% of Ex-work cost per annum will be considered for all bidders for 8 years after warranty.
Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted.
1) | If 2 years warranty and AMC 2% of ex works cost per annum will be considered for all bidders for 8 years after warranty . conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2) | Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3) | Only Yes or No should be mentioned in compliance report and additional information should in remarks column.