## Sheet1
Serial # 4
TMH/TMH/2022-23/CAP/PT/0010
TATA MEMORIAL HOSPITAL
PURCHASE DEPARTMENT
Department: Nursing, TMH
Technical Specifications : Cytotoxic Cabinet (Qty: 04 nos)
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Probable end of support:
1 | Should be a cytotoxic cabinet of Class II, Type A2 with approx. 30% exhaust and 70% recirculation air-flow specifically required for use in preperation of cytotoxic medicines and toxic reagents, etc.
2 | Should have specific system or filters that should protect the operator from cytotoxic substances i.e safe filter change procedure should be available.
3 | Size : 4 feet.
4 | Cabinet Construction:
a) | ~ Main body: 18 gauge electrogalvanised steel, with epoxy powder coated finish.
b) | ~ Work zone: 16 gauge stainless steel, type 304 with 4B finish.
c) | ~ Side walls: 16 gauge, SS 304, 4B finish.
5 | Filters -
a) | ~ Downflow: H14 grade HEPA filter gives 99.97% down to 0.3 microns
b) | ~ Exhaust: H14 type HEPA filter gives 99.97% down to 0.3 microns
6 | Should have LCD display indicating airflow velocity, alarms, UV lights functioning and other imp functions.
7 | The front of the cabinet should be angled so as to help minimize glare on window to the user and integrated arm rest.
8 | Should have UV light for decontamination with on/off timer.
9 | Should be certified to NSF 49 or EN12469
10 | HEPA filters should be EN-1822 certified.
11 | Should be CE/FDA/BIS/ICMED approved.
12 | Cabinet should be supplied with support floor-stand with caster wheels.
13 | Should have built in 5/15A electrical socket.
14 | Should be supplied with 01 no of gas cocks at side walls.
15 | Internal fluorescent lighting Fixture should be such that light should not be facing the user and it should be out of the air stream.
16 | Should display filter life to be prepared with a set of spare filters.
17 | Should be supplied with standard make UPS with atleast 30 mins backup.
18 | Should perform yearly calibration/Validation and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABL/NABH norms). To submit SOPs for PMS at the time of installation/commissioning.
19 | Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
20 | Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.(Quote Rate)
21 | Please specify pre installation requirements [electrical, HVAC, Compressed air ( please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour) ; if any of the above is essential]
22 | Please specify footprint size & weight.
23 | Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
24 | Predispatch inspection at factory will be required if desired by authorities.
25 | Please visit the site to check whether quoted model can be accomodated in the available space.
26 | Should be CE/FDA/BIS/ICMED approved
27 | Please specify power consumption.
28 | Unpacking and shifting the consignment to the installation site is to be included in the scope of supply. Bidder/manufacturer/authorized service provider should take responsibility to lift /shift the consignment from unloading site to the installtion site. Unloading site shall be Stores department, Main Building Basement,TMH. If needed ,bidder has to arrange for the laboures at no charge to TMH.(Before submitting the quotation,bidders may visist TMH to Know unloading and installtion site.
29 | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
30 | Response Time: Should not be more than 06 hrs from lodging a breakdown complaint on toll free or by email.
31 | Warranty- 2 years
32 | CMC- 5% of Ex-works cost per annum will be considered for all bidders for 8 years after warranty
33 | Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted.
34 | Buyback: Against old Cytoxic Cabinet Dyna make. Qty: 04 Nos
1) | If 2 years warranty and CMC maximum 5% of Ex works cost per annum will be consided for 8 years after warranty . conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2) | Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3) | Only Yes or No should be mentioned in compliance report and additional information should in remarks column.