## Sheet1
Serial # 4
TMH/VIZ/2022-23/CAP/PT/0013
TATA MEMORIAL HOSPITAL
PURCHASE DEPARTMENT
DEPARTMENT: Dharmshala, Brachytherapy OT, Nursing and ICU, HBCHRC Vizag
Technical Specifications of " Syringe Pump" Qty: 57 Nos (Qty: 20 Nos for Dharmshala, Qty: 03 Nos for Brachytherapy, Qty: 08 Nos for Nursing and Qty: 26 No for ICU, HBCHRC Vizag)
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of item | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Probable year of end of support:
1 | Syringe pump should accept 20 and 50 ml syringes of all standard companies.
2 | Should have flow rate range from 0.1 to 500 ml/hr depending on syringe capacity.
3 | Should have purge and bolus facility.
4 | Should have occlusion pressure detection.
5 | Should have rechargeable Battery with backup of atleast 4 hours from fully charged
6 | Battery capacity indicator.
7 | Should have various Alarms such as near empty, end, occlusion, low battery, etc.
8 | Volume Infused Display
9 | Should be portable and easy to carry.
10 | Should be provided with pole clamp to fix on IV stand.
List of all optional accessories & consumbales to be provided and quote rate for the same.
Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABH norms). To submit SOPs for PMS at the time of installation/ comissioning.
Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
Please specify pre installation requirements (electrical, HVAC compressed air (Please specify rquired air pressure), Water requirement (RO/DI/ Distilled water with its pressure, flow rate per hour); If any of the above is essential)
Please specify footprint size & weight
Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
Plesae specify power consumption.
Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
Response Time: Should not be more than 12hrs from lodging a breakdown complaint on toll free or by email.
Warranty: 2 years
AMC: AMC @ maximum 2% per annum for 8 years on Ex-works cost after warranty.
Back to back assurance to be taken by the supplier from the OEM to supply spares for minimum 10 years & to be submitted.
1. If 2 years warranty, AMC @ maximum 2% per annum for 8 years on Ex-works cost after warranty. and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3. Only Yes or No should be mentioned in compliance report and additional information in remarks column.