## Final
Serial # 4
TMH/TMC/2022-23/CAP/PT/0014
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
DEPARTMENT: ICU, HBCHRC Vizag & HBCHRC Mullanpur
Technical Specifications of " Dialysis machine" QTY- 03 Nos (Qty: 02 nos for HBCHRC Vizag & Qty: 01 no for HBCHRC Mullanpur)
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
SN. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Probable end of support:
1 | Should have facility for conventional ,High flux dialysis.
2 | Machine should have inbuilt online dialysate fluid filter system for ultrapure dialysate delivery.
3 | Battery backup of 20 minutes atleast to run complete machine.
4 | Should have Na, Bicarbonate and UF profiling
5 | Dialysate temperatures selectable between 35 degrees Centigrade to 39 degrees Centigrade or better .
6 | Variable conductivity setting between 13 to 15 ms/cm or better .
7 | Should have variable dialysate flow 300-800 ml/min or better.
8 | Should have facility to show trends curve of all parameters.
9 | Heparin pump with syringe of different sizes
10 | Ultrafiltration 0.1 to 3 litres/hr or better. The in and out fluid circuit must be separated so that there is no chance of contamination in the event of membrane rupture.
11 | Treatment parameter should be displayed on monitor.
12 | Should have integrated heat and chemical disinfection facility with one standard time.
13 | Should have accurate feedback control conductivity mixing technique.
14 | Should have drain facility.
15 | Should have accurate Ultrafiltration control by flow by volume control measurement technique.
16 | All important data should be presetted so that machine can be used anytime without feeding data every time.
17 | Should have automatic self test facility.
18 | Should have soft touch keypad with large colour TFT Screen/touch screen.
19 | Automatic diagnosis of malfunctioning with on line ability to show the faults.
20 | Blood pump rate from 30 -500 ml/min or better.
21 | Audio visual alarms on limit violation of conductivity,blood leak, air leak, transmembrane pressure alarms,, Dialysis temperature alarm,dialysis can empty alarm, end of disinfection alarm,bypass alarm and blood pump stop alarm.
22 | Should have inbuilt NIBP.
23 | All consumables required for installation and standardization of system to be given free of cost.
Cost of consumables should be quoted seperately with breakup.
The unit should comply with standard and should have CE or US FDA marking/ BIS/ ICMED approved.
Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABL norms)
Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
Please specify pre installation requirements (electrical, HVAC etc.)
Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
Response Time: Should not be more than 12hrs from lodging a breakdown complaint on toll free or by email.
Warranty: 2 years
CMC: Maximum 5% of Ex works cost per annum will be consideed for 8 years after warranty
Back to back assurance to be taken by the supplier from the OEM to supply spares for minimum 10 years and to be submitted .
1) | If 2 years warranty and CMC Maximum 5% of Ex works cost per annum will be consideed for 8 years after warranty . conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2) | Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3) | Only Yes or No should be mentioned in compliance report and additional information should in remarks column.