## Sheet1
Serial # 4
TMH/TMC/2022-23/CAP/PT/0015
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Dept Name: Nursing, HBCHRC Vizag & HBCHRC Mullanpur
Technical Specifications: Suction Machine (Portable) Qty: 66 nos (Qty: 19 Nos for Vizag & Qty: 47 Nos for Mullanpur)
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Year of Probable end of support:
1 | Should be designed for draining blood and other fragmetic secretions in operation room and emergency room in hospital.
2 | Should be fitted with oil immersed noiseless motorized vacuum pump.
3 | Cabinet made from stainless steel.
4 | Two glass jars on the top of having minimum capacity of 2ltrs fitted with rubber air tight lids and overflow safety device.
5 | Should be supplied with adequately long pressure tubing providing required pressure.
6 | Should have vacuum control by knob.
7 | Should be mounted on four caster wheels.
8 | Should have motor of 1/4Hp capacity and power consumption not more than 250 Watts
9 | Should have vacuum at least between-100mmHg to at least 575 mmHg (-75 kPa).
Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABL norms). To submit SOPs for PMS at the time of installation/commissioning.
Performance certificates of IQ,OQ,PQ tests to be submitted at the time of installation/commissioning ; if desired by the end user. (Optional)
Should be supplied with standard make UPS with atleast 30 mins backup. (Optional)
Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
Please specify pre installation requirements [electrical, HVAC, Compressed air ( please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour) ; if any of the above is essential]
Please specify footprint size & weight.
Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
Predispatch inspection at factory will be required if desired by authorities.
Please visit the site to check whether quoted model can be accomodated in the available space.
Please specify power consumption.
Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
Warranty- 2 years
AMC- 2% of exworks cost per annum will be considered for all bidders for 8 years after warranty
1) | If 2 years warranty and AMC 2% of ex works cost per annum will be considered for all bidders for 8 years after warranty . conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2) | Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3) | Only Yes or No should be mentioned in compliance report and additional information should in remarks column.