## Sheet1
Serial # 4
TMH/BHU/2021-22/CAP/PT/0105
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Dept: Department of Surgical Oncology Dept, MPMMCC Varanasi
Technical Specifications of "Rhino-Laryngo Video Cart System" Qty-2
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at HBCH/MPMMCC (Tata Memorial Centre), Varanasi.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Name of the Make:
Model:
Year of Introduction:
Probable end of support:
A. | Full High-Definition System should consist of following items:
Note: | Only one scope with one cart (i.e Adult scope with one cart system and Pediatric scope with another cart system)
1 | A Chip on Tip Rhino-Laryngo Videoscope (Adult)
2 | A Chip on Tip Rhino-Laryngo Videoscope (pediatrics)
3 | Full High-Definition Video Processor system
4 | Powerful LED light source
5 | 26” or more Full High-definition medical grade monitor
6 | Trolley to accommodate all equipment
7 | Laptop With a recording Software.
8 | Printer
B. | Rhino- Laryngo videoscope (Adult) should have :
1 | A 2–4-way angulation of scope.
2 | Four or more no. of remote-control switches on control body.
3 | Compatible with electrosurgical and laser treatments
4 | Distal end & Insertion tube outer diameter should be 5 mm or less
5 | Field of view should be 85 degree or more
6 | Depth of field should be 2 mm-50mm
7 | Angulation range should be approx. Up 130 deg & Down 130 deg
8 | Working length should be around 300mm or more
9 | Instrument Channel should be around 1.5 – 2.5 mm
10 | Should be supplied with compatible Leakage Tester
C. | Rhino-Laryngo Videoscope (Pediatric) should have:
1 | A 2–4-way angulation of scope.
2 | Four or more no. of remote-control switches on control body
3 | Distal end & Insertion tube outer diameter should be 2.9 mm or less
4 | Field of view should be 90 degree or more
5 | Depth of field should be 3.5 mm-50mm
6 | Angulation range should be approx. Up 130 deg & Down 130 deg
7 | Working length should be around 300 – 350 mm
8 | Should be supplied with compatible Leakage Tester
D. | Full High-Definition Video Processor system should have:
1 | A full high-definition processor should have resolution of 1920 * 1080 pixels or better
2 | Should have compatibility for selecting Progressive / Interlaced output
3 | Optical image enhancement of capillary vessels and fine patterns in the superficial layer of mucosa for early detection of lesions
4 | Should have a USB slot so as to take still pictures of Endoscope images.
5 | Compatible with Analog and Digital (Y/C, RGB, VBS COMPOSITE, HD-SDI, SD-SDI, DVI) to reproduce high definition images / video.
6 | Should have automatic white balance
7 | Should be able to provide 16:9 and 16:10 output for a HDTV monitor and should be compatible with Analog and digital (HD-SDI and DVI output) to reproduce high-definition images/video.
E. | Powerful Led Source (equivalent to 300W Xenon light source) should have:
1 | A Powerful led light Source having an illumination of equivalent to 300W xenon light source
2 | Automatically adjusts light intensity to achieve ideal illumination
3 | The life of light source should be 10000 hrs. or more
4 | It is better if stroboscope compatibil light source is available or can be connected in future
F. | 26” or more Full High-definition medical grade monitor should have:
1 | Should have minimum 26 inches size
2 | A medical grade high-definition monitor with resolution of minimum 1920*1080 or better
3 | Should have 16:10 aspect ration or better
4 | LCD Screen.
5 | Should have PIP facility.
6 | Should be from same manufacturer.
G. | Trolley:
1 | Trolley should accommodate all Equipment.
2 | Should have isolation transformer.
H. | LAPTOP with Recording Software
1 | A Suitable laptop should be provided.
2 | It should have a Recording Software to record the procedure.
3 | Recording software should be User friendly.
I. | PRINTER
1 | A color Printer should be provided.
2 | Printer should be from a reputed company
NOTE :- All the Equipment’s should be From Same manufacturer Except LAPTOP, Recording Software & Printer.
J. | Standard points :
1 | Should have USFDA & European CE.
2 | Power supply- 100-240VAC, 50-60Hz.
3 | Should be able to withstand temperature and humidity fluctuations in case of failure of air.
4 | Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
5 | Back to back assurance to be taken by the supplier from the OEM to supply spares and local service support till the tenure of 12 years from the date of installation
6 | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
7 | Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
8 | Demo of the quoted model (Anywhere across the globe or country), will be mandatory at the cost of bidder if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
9 | If more then one model is offered, then they are to be separately filled and submitted according to the given format.
10 | A list of installations with the address and contact numbers to be provided. (User list should be for the quoted model)
11 | Warranty: 2 years
12 | AMC and CMC : AMC 2% of ex works cost per annum will be considered for all bidders for 8 years after warranty and CMC 5% of ex works cost per annum will be considered for all bidders for 8 years after warranty (Quote separately)
* | Only Yes or No should be mentioned in compliance report and additional information should be in remarks column.
* | If any of the above conditions wherever applicable are not adhered to, TMC reserves right to reject the offer even after finalizing the contract.
* | Please mention whether UPS/Stabilizer/R.O is required for the equipment. If so should be included in the offer.
* | TMC has the right to ask for the replacement of machine or reject the contract if more then 10 breakdown will happen in the machine during first three months.
* | Should provide all the related documents for the verification of compliance that will be filled in respect of quoted model (example Catalogue of mentioned model)
If 2 years warranty. AMC 2% of ex works cost per annum will be considered for all bidders for 8 years after warranty and CMC 5% of ex works cost per annum will be considered for all bidders for 8 years after warranty . conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.