## Sheet1
TMH/TMC/2021-22/CAP/PT/0106
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Department : Surgery , Varanasi & Mullanpur
Technical specifications - Operation Table Mid End, Qty: 05 Nos (Qty: 04 Nos for MPMMCC Varanasi and Qty: 01 no for HBCHRC Mullanpur)
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation.
| Description of Item | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Year of Probable end of support:
1 | It should be a mobile universal electro-hydraulic driven operating table for performing all Minor and Major onco surgeries .
2 | All below functions of the table should be electro-hydraulic driven & adjustable via corded hand control & override panel for easy access.
a | Height adjustment
b | Lateral tilt
c | Trendelenburg & Reverse trendelenburg
d | Back up/down movement
All below functions of the table should be manually adjustable:
e | Leg plate up/down movement
f | Lock /unlock
g | Head Rest up/down movement
3 | General operating table features should be:
a | It should have central foot pedal for locking or base locking feature.
b | Base cover made of rugged glass fiber reinforced laminate/ stainless steel,resistant to impact, breakage and disinfectants,dark colored dyed throughout with additional scratch-proof enamel finish.
c | Entire table top frame and column cladding should be made of stainless steel to avoid rusting.
d | Special double layer soft padding with latex free bi-elastic cover , fixed with velcro straps on plastic carrier plates, radio translucent.
e | The table top is fully free without transverse shadows on X-ray images with guide rails / alternate mechanism for X-ray cassettes inserted from the head end up to and including the pelvis.
f | Patient orientation should be possible on both sides of the table top.
g | Minimum loading capacity should be 250Kg
4 | Electrical Specifications should be:
a | Special-design, maintenance-free rechargeable batteries with capacity for about a week's use/Minimum 20 Operations on single charge,in the operating room.
b | Recharging of the batteries and supply of the operating table by means of mains cord.
c | Nominal mains voltage : 230-240V AC via a mains cord
d | Mains frequency , 50/60HZ.
5 | Technical specifications should be :
a | Height without padding : 700-950mm or better
b | Trendelenburge / Reverse Trendelenburg: +25°/-25° or better
c | Lateral Tilt , both directions : 15° or better
d | Back plate UP/Down :70° / -40° or better
6 | Standard Accessories to be given for each table:
1 | Head Rest with special foam padding
2 | Adapters pair for head rest
3 | Pair of Arm boards with pad and clamps
4 | Anesthesia screen with clamp
5 | Body Strap -2nos with each table
6 | Pair of Abductable Leg Plates with side rails
7 | Backlit cable connected hand control
8 | Head ring gel
General terms:
Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABL norms). To submit SOPs for PMS at the time of installation/commissioning.
Performance certificates of IQ,OQ,PQ tests to be submitted at the time of installation/commissioning ; if desired by the end user. (Optional)
Should be supplied with standard make UPS with atleast 30 mins backup. (Optional)
Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
Please specify pre installation requirements [electrical, HVAC, Compressed air ( please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour) ; if any of the above is essential]
Please specify footprint size & weight.
Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
Predispatch inspection at factory will be required if desired by authorities.
Please visit the site to check whether quoted model can be accomodated in the available space.
Please specify power consumption.
Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
Response Time: Should not be more than 06 hrs from lodging a breakdown complaint on toll free or by email.
Warranty- 2 years
CMC- 5% of Ex-works cost per annum will be considered for all bidders for 8 years after warranty
Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted.
1) If 2 years warranty, CMC @maximum 5% per annum of ex-works cost of the equipment will be considered for all bidders for 8 years after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject offer.
2) Please mention whether UPS / Stabilizer is required for the equipment. If so, should be included in the offer.
3) Only yes or no should be mentioned in the compliance report and additional info in remarks column.