## Sheet2
Serial # 4
TMH/MUL/2021-22/CAP/PT/0109
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Dept Name: Surgery, HBCHRC, Mullanpur
Technical specifications - "CO2 Laser machine" QTY: 1 NO
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Year of Probable end of support:
A) | Laser System:
1 | LASER SYSTEM SHOULD BE A CARBON DIOXIDE LASER WITH A WAVELENGTH 10.60 MICRO METERS, INFRARED
2 | LASER SYSTEM SHOULD BE A CARBON DIOXIDE LASER WITH A WAVELENGTH 10.60 MICRO METERS, INFRARED
3 | LASER MACHINE SHOULD HAVE POWER OUTPUT 1-40 WATTS.
4 | IT SHOULD HAVE 5Mw RED DIODE AIMING BEAM,635NM, ADJUSTABLE INTENSITY
5 | THE BEAM DELIVERY SHOULD BE THROUGH (BOTH 1 & 2) 1) A 7-JOINT, FIXED MIRROR, SPRING BALANCED ARM, THE REACH OF THE ARM SHOULD BE AT LEAST 120 CM WITH 360 DEG ROTATION. 2) A LIGHT WEIGHT CARBON DIOXIDE GLASS HOLLOW FIBER. CO2 FIBER SHOULD BE 2 METER LONG, 1.04MM OUTSIDE DIAMETER, STERILE, SINGLE / MULTIPLE USE, 2.0 METER LONG.
6 | IT SHOULD BE EQUIPPED WITH ONE TOUCH TAB/SWITCH TO CHOOSE EITHER WAVE GUIDE OR ARTICULATED ARM MODALITY WITHOUT CHANGING ANY PART.
7 | SPOT SIZE: 295ΜM AT FIBER OUTPUT. UP TO 40 WATT
8 | IT SHOULD BE MICROPROCESSOR BASED.
9 | IT SHOULD HAVE A SEALED CO2 LASER TUBE.
10 | IT SHOULD HAVE CONTINUOUS, SINGLE PULSE AND REPEAT PULSE TISSUE EXPOSURE MODES.
11 | IT SHOULD HAVE CONTINUOUS POWER (CW) OF 01 – 40 WATTS
12 | IT SHOULD HAVE A SUPER PULSE POWER OF 0.5 – 15 WATTS.
13 | • IT SHOULD HAVE A TIMED EXPOSURE OF FOLLOWING DURATIONS: 1) ON TIME (SINGLE PULSE) – 0.05 – 1.0 SEC. AT 1.0 TO 4.5 WATTS - 0.01- 1.0 SEC AT 5-40 WATTS 2) ON TIME (REPEAT PULSE) – 0.05 – 1.0 SEC AT 1- 4.5 WATTS - 0.01 – 1.0 SEC AT 5-40 WATTS
14 | IT SHOULD HAVE A REPEAT DELAY, OFF TIME, 0.01 TO 1.0 SEC.
15 | IT SHOULD HAVE AT LEAST 100 USER DEFINED MEMORY SETTINGS.
16 | IT SHOULD HAVE A 0.2MM FOCUSED HAND PIECE.
17 | IT SHOULD HAVE AT LEAST TWO BACTERIAL FILTERS.
18 | IT SHOULD HAVE THREE LASER SAFETY GLASSES.
19 | IT SHOULD HAVE A SELF CONTAINED CLOSED LOOP COOLING SYSTEM.
20 | IT SHOULD HAVE A MULTI –COLOUR TOUCH SCREEN PANEL
21 | IT SHOULD BE EQUIPPED WITH INTEGRATED ANIMATED ACCESSORIES VIDEOS DEMONSTRATING HOW TO SET UP IT BEFORE STARTING APPLICATION/SURGERY.
22 | IT SHOULD HAVE A USER FRIENDLY GRAPHIC DISPLAY TO PROVIDE STEP BY STEP OPERATING INSTRUCTIONS.
23 | IT SHOULD BE COMPATIBLE WITH 230V, 3A, 50HZ POWER SUPPLY
B) | MICROMANUPLATOR WITH FOLLOWING REQUIREMENTS FOR MICROLARYNGEAL LASER SURGERY:
1 | IT SHOULD HAVE AN OPTICAL DESIGN TO ASSURE PERFECT CO-INCIDENCE OF THE DIODE AND CO2 BEAMS EVEN AT HIGHEST MICROSURGICAL MAGNIFICATIONS.
2 | IT SHOULD BE EASILY ADJUSTABLE AND SHOULD HAVE VARIABLE WORKING DISTANCE FROM 200MM TO 400MM.
3 | IT SHOULD HAVE CONTINUOUSLY VARIABLE DEFOCUS WITH A USER ADJUSTABLE DEFOCUS LIMITER.
4 | ITS JOYSTICK HANDLE SHOULD BE TENSION ADJUSTABLE AND AUTOCLAVABLE.
5 | IT SHOULD BE USER SELECTABLE FOR LEFT OR RIGHT HAND CONTROLS.
6 | IT SHOULD BE LIGHTWEIGHT, TO MAINTAIN BALANCE OF THE SURGICAL MICROSCOPE
7 | IT SHOULD HAVE A MINIMUM SPOT SIZE OF 160 MICRONS.
8 | IT SHOULD HAVE A FOCUS RANGE OF 0.16 MM – 0.27 MM.
9 | IT SHOULD HAVE MAXIMUM DEFOCUS RANGE OF 2.8 MM– 4.6 MM.
10 | IT SHOULD HAVE A POWER TRANSMISSION OF GREATER THAN 90%, WITH UNLIMITED POWER INPUT.
11 | IT SHOULD HAVE A ROBOTIC LASER MICROSURGERY SYSTEM WITH FOLLOWING REQUIREMENTS:
12 | IT SHOULD HAVE BEAM SCAN SHAPE: LINEAR & CURVED INCISIONS: 0.3MM TO 5.0 MM IN LENGTH (USER DEFINED), 0.7 MM TO 3 MM FOR PAPILLOMATOSIS.
13 | IT SHOULD HAVE A PENETRATION DEPTH OF 0.2 MM TO 2MM. (USER DEFINED)
C) | ORAL, PHARYNGEAL AND NASAL HANDPIECE SET WHICH INCLUDES
1 | 230MM HANDPIECE UNIT (CVD OPTICAL UNIT, PORTS HOLDER,M CONICAL MAIN EXTENDER, CONTAMINATION COLLECTOR), EXTRA CONICAL MAUN EXTENDER, BACKSTOP EXTENDER – 3 NOS, TIP EXTENDER – 3 NOS, STRAIGHT TIP, KAMAMI NASAL TIP – 3 NOS, KAMAMI TONSIL TIP – 3 NOS, 90 DEGREE ANGLED MIRROR TIP EXTENDER, CLEANING BRUSH, TYGON TUBE(8MM ID,1.5M LONF) W/ REDUCER FITTING.
D) | FIBER ACCESSORIES
1 | REUSABLE CO2 FIBER- 01 NOS
2 | RIGID HAND PIECE KIT AT LEAST 8 RIGID HAND PIECES
3 | 60MM, STRAIGHT, STRAIGHT TIP, 180MM, STRAIGHT, STRAIGHT TIP, 60MM, STRAIGHT, CURVED TIP, 140MM, STRAIGHT, CURVED TIP 180MM, STRAIGHT, CURVED TIP, 240MM, BENT, CURVED TIP, 140MM, BENT, STRAIGHT TIP, 240MM, BENT, STRAIGHT TIP.
4 | ENDOSCOPE PROTECTION SHEATH – 2 NOS LENGTH: 640 MM, OD: 1.7 MM
5 | HAND PIECE BENDING TOOL
6 | HAND PIECE CLEANING KIT: INCLUDES 3 CLEANING BRUSHES AND 20 EXTRA SILICONE TUBE FOR HANDPIECES
7 | BENDING AND CUTTING TOOLS TO REUSE FIBER
8 | STERILIZATION TRAY FOR FIBERS.
Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABH norms)
Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
Please specify pre installation requirements (electrical, HVAC etc.)
Please specify footprint size & weight.
Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
Please visit the site to check whether quoted model can be accomodated in the available space.
Please specify power consumption.
The unit should be CE/FDA/ICMED/ BIS approved.
Predispatch inspection at the factory will be required it desired by authorities
Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
Warranty: 2 years
CMC: CMC maximum 5% of Ex works cost per annum will be consideed for 8 years after warranty
Back to back assurance to be taken by the supplier from the OEM to supply spares for minimum 10 years and to be submitted.
1) | If 2 years warranty and CMC maximum 5% of Ex works cost per annum will be consideed for 8 years after warranty . conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2) | Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3) | Only Yes or No should be mentioned in compliance report and additional information should in remarks column.