## Sheet1
Serial # 4
TMH/TMC/2021-22/CAP/PT/0136
TATA MEMORIAL HOSPITAL
PURCHASE DEPARTMENT
Department - Transfusion Medicine, HBCHRC Vizag & HBCHRC, Mullanpur
Technical specifications of "Automated Hemoglobinometer(Reagent Rental contract for 3 years plus 1 year extendable)", Quantity 06 nos. (Qty: 04 Nos for HBHRC Vizag & Qty: 02 Nos for HBCHRC Mullanpur)
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quote model should be latest introduced with end of life not beore 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Probable end of support:
1 | Should be micro-processor based battery operated.
2 | Accuracy should not be less than ± 1.5% or ± 0.5g/ dl compared to the international reference method for hemoglobin (ICSH method)
3 | Should be based on photometric chemical free method
4 | Analyser should be factiry calibrated and calibration is verified every time when analzer turns on or should provide standard control/ Cuvettes for daily calibration or QC check.
5 | should require capillary, arterial or venous blood Sample.
6 | Should display the result within 15-60 seconds depending upon the hemoglobin concentration/ turbidity of the blood and measures in g/dl
7 | Should have dual wavelength operation for turbidity compensation.
8 | Should perform quality control/ calibration of equipment at user defined intervals.
Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards (As per NABL norms.) To submit SOPs for PMS at the time of installation/ Commissioing.
Performance certificates of IQ,OQ,PQ test to be submitted at the time of installation/commissioning; if desired by the end user. (Optional)
Should be supplied with satandard make UPS with atlest 30 mins breakup (Optional)
Please submit the technical data sheet along with the brochure as a tender document with referance to the quoted make & model.
Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
Please specify the pre installation requirements (Electerical, HVAC, Compressed air (Please specify required air pressure), water requirement (RO/DI/Distilled water with its pressuer, flow rate per hour); if any of the above is essential)
Please specify footprint size & weight.
Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
Predispatch inspection at factory will be required if desired by authorities.
Please visit the site to check whether quoted model can be accomodated in the available space.
10 | Please specify power consumption.
12 | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
13 | Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
CRITERIA FOR RR:
Venodr to offer cost per test
1) Please mention whether UPS / Stabilizer is required for the equipment. If so, should be included in the offer.
2) Only yes or no should be mentioned in the compliance report and additional info in remarks column.