## Sheet2
Serial # 4
TMH/MUL/2021-22/CAP/PT/00137
TATA MEMORIAL HOSPITAL
(Purchase Department)
Dept- Pathology, HBCHRC Mullanpur
Technical Specification- Cassette Labeller, Qty- 1 no
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr no | Item Description | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Probable end of support:
Specification:
1 | LASER or Ink jet technology should be used to mark cassettes with high resolution text which can be readily scanned from sample receiving till archiving.
2 | Should produce durable text which should withstand harsh reagents without fading text for minimum of 10 years
3 | Should have high printing resolution of 300 dpi or more with alphanumeric characters, one dimensional codes, data matrix, logos.
4 | Should be able to print minimum of 10 cassettes or more per minute
5 | Should be able to interface with existing tracking system.
6 | Software should be user friendly to design and edit cassette printing format with the possibility of using extensive and standardized protocols.
7 | Software should support linear bar code, 2D bar code, text, special characters and fonts
8 | Should have atleast 6-8 magazines with around 80 cassettes per magazine
9 | Instrument should indicate status for 'reset', 'power', 'online', 'error', 'magazine loading status', 'Ink status'.
10 | Should have network connection LAN, USB and should be provided with unloading station.
11 | Should be provided with Touch PC with atleast 24'' monitor , barcode scanner USB kit with scanner stand.
12 | Should have STAT/CITO printing for manual feeding.
13 | Use of all types of cassettes with write-able surface angles (25°- 45°)
14 | Power supply- 100-240VAC, 50-60Hz
15 | Should be FDA approved or CE marked/ BIS/ ICMED
General Terms:
1 | Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABH norms)
2 | Performance certificates to be submitted at the time of installation / commissioning.
3 | Should be supplied with standards make UPS with atleast 30 Min backup.
4 | Submit the technical data sheet along with the brochuer as a tender document with reference to the quoted make & model.
5 | Please specify list & cost of consumables/ consumable spares (i.e spares need to be replaced at regular intervals, may be quarterly/half yearly/ yearly such as annual maintenance kit etc.) if any.
6 | Specify pre installation requirements [electrical, HVAC, Compressed air ( please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour) ; if any of the above is essential]
7 | Specify footprint size & weight.
8 | Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
9 | Predispatch inspection at factory will be required if desired by authorities.
10 | Visit the site to check whether quoted model can be accomodated in the available space.
11 | Please specify power consumption.
12 | Local Service Support: Should have local service support/service engineer for attending the breakdown calls.
13 | Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
Warranty: 2 years
CMC: CMC @ maximum 5% per annum for 8 years on Ex-works cost after warranty.
Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted.
1. If 2 years warranty, CMC @ maximum 5% per annum for 8 years on Ex-works cost after warranty. and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3. Only Yes or No should be mentioned in compliance report and additional information in remarks column.