## Sheet1
Serial # 4
TMH/BHU/2021-22/CAP/PT/0054
TATA MEMORIAL HOSPITAL
PURCHASE DEPARTMENT
Department - Aneastheisa Dept , MPMMCC Varanasi.
Technical Specification of "Advanced Anaesthesia Delivery System", Quantity 4 nos.
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model:
Make:
Year of Introduction:
Year of Probable end of support:
1 | Durable with a sturdy kit.
2 | Antistatic wheels with brakes
3 | Shelves for monitors
4 | The anesthesia workstation should have in built ventilator with min 12 inch TFT display.
5 | The anesthesia workstation should have electronic gas mixing siutable for adult, children & neanates.
6 | The anesthesia workstation should be suitable for low and minimal flow anesthesia application.
7 | The anesthesia workstation should have automatic calculations and presetting of patient specific ventilation settings via ideal body weight.
8 | It should have fully automatic menu drive self test and user check list. In case of emergency it should be possible to bypass the self test.
9 | Gas delivery system:
a | Unit should have primary connection for central gas supply for Air, O2 & N2O and as a backup machine should have provision for connecting oxygen and nitrous pin index cylinders.
b | Machine should have elecronic fresh gas delivery & monitoring system with automatic switchover to air in case of failure of oxygen.
c | Audible and visual alarms for airway pressure, disconection, oxygen failure.
d | Anti-hypoxia safety mechanism to ensure that O2 concentration doesn’t fall below minimum 23% at any time.
e | O2 safety flow adjustable from 0 to 10L/min or more, runnig through the vaporiser for emergency backup use.
f | Fresh gas fow setting from 150ml /min(or less) to 12L/min
g | Electronic flowmeter for fresh gas flow.
h | Provision of additional O2 port for powering a ventury device.
i | Machine should also have an independent fresh gas outlet for connection to bain's or magill circuit.
10 | Breathing system:
a | Compact breathing system suitable for minimal flow anaesthesia with least patient circuit volume including absorber etc approx 3L (excluding bag) or less for fast response to change in fresh gas composition.
b | Fresh gas decoupled / compensated breathing system for adult & children.
c | should have mechanism to prevent moisture condunsation (eg. Heating circuit or equivalent )
11 | Ventilator:
a | Ventilators for controlled ventilation of adults as well as children
b | Electronically controlled electrically or pneumatically driven ventilator.
c | Automatic breathing circuit compliance correction.
d | Provision for manual and mechanical ventilation
e | IPPV or equivalent with ability to adjust plateau time
f | pressure controlled ventilation.
g | SIMV with PCV & VCV
h | Pressure support, PS with CPAP, PS with PCV&VCV,SIMV with minimum Tidal volume upto 5ml.
i | High peak inspiratory flow upto 120LPM or more
j | Tidal volume delivery from 20 to 1400 ml or more
k | PEEP setting: upto 20 cmH20 or more
l | Respiration rate: 4(or less)-100 bpm
m | I:E : Max 2:1 to 1:8 or better
12 | Ventilator monitoring:
a | Monitoring of volume, pressure & oxygen
b | Should monitor patient compliance
c | should display curve & loops with fast analysis of change in lung mechanism provided by simultaneous display ETCO2,O2 as well as anaesthetic agent and automatic identification of agent .
d | Bar graph display of volumeter , tidal volume and the difference of inspiratory and expiratory tidal volume.
e | Trending for alarms and measured values with graphic trend display.
f | The machine should calculate agent cosumption on a case by case basis and display of fresh gas consuption in the unit logbook
13 | Alarms:
a | The machine should have adjustable alarm limits for all the parameters with auto set alarm function.
b | The machine should have automatic display of MAC values
c | Airway pressure low limit should automatically changes with the PEEP settings changes.
d | Machine should have RS232 connectivity port for interface to patient monitor/ HIS for automatic data acquisition.
14 | Vaporisers:
a | temperature /pressure compensated and flow independent vaporizer.
b | Vaporizer should have extended delivery range from 0-6% volume
c | Vaporizer should have transport lock to provide hermetic sealing of agent chamber during transport & storage.
d | The vaporizer design should be maintenance free.
15 | Agent Monitoring:
a | should display agent analysis, N2O & MAC value with color coding of agents.
16 | Accessories
a | Should provide all standard acceossories
b | Open and closed circuit for both adult and pediatric
c | List of optional attachemts and accessories compatible with the specified model.
17 | Should work on mains Power and with battery backup of 45 minutes minimum with fully charged batteries
18 | Should be able to withstand temperature and humidity fluctuations in case of failure of air conditioning.
19 | The unit should comply with standards and should have CE/FDA/BIS/ICMED Approval.
20 | Warranty : 2 years
21 | CMC : Maximum 5% per annum for 8 years on Ex-works cost after warranty. (Included 2 preventive maintenance/ Year and all breakdown visits. All bought items should also be included in the warrenty. CMC should include upgrades / updates cost, No separate software licence fees will be payable.)
22 | Back to back assurance to be taken by the supplier from the OEM to supply spares for minimum 10 years and to be submitted to TMH.
23 | Demo of the quoted model (Anywhere across the globe), will be mandatory at the cost of bidder if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
1. If 2 years warranty,CMC maximum 5% per annum for 8 years on Ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3. Only Yes or No should be mentioned in compliance report and additional information should in remarks column.