## Sheet1
Serial # 4
TMH/VIZ/2021-22/CAP/PT/0055
TATA MEMORIAL HOSPITAL
PURCHASE DEPARTMENT
Department - Dental, HBCHRC Vizag.
Technical specification of "Dental Chair", Quantity 02 nos.
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model:
Make:
Year of Introduction:
Probable end of support:
A | Dental Chair : Quantity 02 nos.
1 | Physiologic body contoured, electro-hydraulically/electro-mechanical driven chair with smooth movement.
2 | Programmable chair (with four or more programs : programs for patient entry, for rinse and for patient exit). These programs can be set with touch of a button.
3 | User friendly single foot control for chair & Unit.
4 | Please specify the position of touch panel control.
5 | Back rest should help operator to move closer to patient & get more knee space below the back rest which enables operator to work in a comfortable position.
6 | Chair should offer two arm rest. Right arm rest should be movable for easy patient entry & exit.
7 | Height range is from 415mm to 720mm or better to improve patients entry & exit.
8 | Back Rest Angle from 2o lower from horizontal to 70o in upright position.
9 | Chair should have 2 axis head rest. It should move up for tall patients and accommodate pediatric patients also.
10 | Iluminance level of LED OP light can be set from the chair touch pad with different intensities.
11 | The body should be made up of cast aluminium alloy for better stability and longevity.
• | Traditional Delivery System with 5 position head
1 | Unit should be over the patient delivery type.
2 | Touch screen on the unit head for convenient operations
3 | Unit holders should be integrated into unit head.
4 | Unit should have following built-in attachments:-
a | Built-in high speed handpiece control (2 nos.)
b | Built-in 1 no. Brushless Micromotor with One Straight & 1 Contra Handpiece 1:1 for normal use should be supplied.
c | Should have preferably one built-in Scaler with 3 scaling tips.
d | Built-in 3 way syringe with optional warm water facility.
5 | All the handpiece & integrated attachments on the unit should be preferably of same manufacturer as Chair & Unit to make warranty claims easy.
6 | Ability to program hand pieces.
7 | Large work table with mat.
8 | All hand pieces tubes should have antibacterial feature to avoid formation of bacterial growth in hand pieces tubing.
9 | Aerosol antisuck mechanism should be demonstrable.
10 | Air Turbine should have 3 air turbine / Air rotor handpieces should be prioritize.
11 | Water spray from all the Handpiece should be controlled individually from the unit.
12 | Water supply to the spittoon should be online from main water line. To facilitate continuous water to the unit without interruption.
• | Cupsidor :
1 | Autoclavable saliva ejector.
2 | Autoclavable High volume evacuator.
3 | Autoclavable 3 way syringe.
4 | Assistance arm is extendable.
5 | Standard multi-function touchpad and designed to be used as a handle for positioning of the assistant’s arm.
6 | Vitreous china bowl with adjustable cup fill sensors for water level and bowl rinse timers.
7 | Cup filler should have auto filling sensor for water level.
8 | Cup filler should preferably have warm water facility.
• | Operating LED Light :
1 | Light head should have three axial movements – Horizontal, Vertical & Optional Axial (diagonal) adjustment.
2 | Hands free control of operating light using the foot switch for ON/OFF.
LED light should have at least 3 intensities. Low, Medium & High.
4 | LED light with colour corrected temperature of 4200 degree Kelvin or better.
5 | Feathered-edge, balanced-intensity light pattern.
6 | Optically designed reflector.
7 | Auto On/Off functions with chair pre-set movements. Or by foot control or by controls on touch pad on doctor’s or assistant’s side
• | Doctor’s Stool
1 | Cast-aluminum base with tile casters
2 | Two-way adjustable lumbar support with Arm Rest.
3 | Integral gas cylinder for height adjustment.
4 | Foot ring for adjusting height of the stool. Conveniently making it hands free adjustment for stool height.
5 | Height range 450mm – 650mm
B | Auxillary Attachments :
• | Apex Locator, Quantity 01 no.
1 | Should utilize the ratio technique of calculating the impedances of 2 frequencies
2 | Should have autocalibration
3 | No zero adjustment should be provided
4 | Should have colour liquid crystal display
5 | Should have Flash bar, Memory bar, Information display, Canal length Indicator Bars
6 | Should be light weight to work on 3-A3 sized dry cells
7 | should have provision of Data Port for connecting it to Endodontic Micromotor to make it integrated Endomotor with Apex Locator if desired
8 | Kit contents should be Apex Locator, 3 pieces of File Holder, 5 pieces of Lip Clip and Tester.
• | Endo Motor, Quantity 01 no.
1 | Cordless endo motor should be integrated with Apex Locator if desired
2 | should have Automatic Functions such as Automatic Torque Reverse, Automatic Torque Slow Down, Automatic Apical Reverse, Automatic Apical Slow Down, Automatic Apical Torque Reduction, Automatic Start and Stop when connected to apex locator.
3 | should have back lit display for programming.
4 | Kit contents should be Endodontic Motor Handpiece with Battery, Charging Unit and Oil.
5 | Data Cable & Contrary Electrode should be supplied.
• | Light Cure Unit, Quantity 01 no.
1 | Should provide uniform light intensity
2 | Should have multiple curing modes to provide optimum output for various cleaning areas
3 | The collimated light should maintain 80% of its intensity at a distance of 5 mm from the light source
4 | Approximate High power should be 1800 mW/cm² or better.
5 | Should be provided with battery charger
6 | Fully charged battery should deliver atleast 400 cycles for 3-second cures or 100 cycles of 10-second cures
C | Motorised Suction Unit, Quantity 02 nos.
1 | Should be used for 1 dental chair
2 | Should have minimal noise approx 53-63dB (A)
3 | Should have current consumption of 2.0 A
4 | Should have electrical power of 0.38 kW
5 | Should have Rotational Speed of 2,850 rpm
6 | Should have suction capacity of 300 litres of air per minute
7 | Should have 100% Duty cycle
8 | Should weight about 12-15kg
9 | Suction unit should come with original cover/housing to reduce the sound.
10 | Suction unit should exhaust filter to prevent venting out of bacteria & viruses.
11 | Filter should be HEPA-H13 filter.
12 | 5 nos. Extra Filters to be supplied with the suction unit.
D | Compressor
Chair Should be compatible for central compressed air line.
13 | Should meet national and international standards ISO/CE/US FDA/BIS/ICMED Certified
14 | Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABL norms). To submit SOPs for PMS at the time of installation/commissioning.
15 | Performance certificates of IQ,OQ,PQ tests to be submitted at the time of installation/commissioning ; if desired by the end user. (Optional)
16 | Should be supplied with standard make UPS with atleast 30 mins backup. (Optional)
17 | Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
18 | Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
19 | Please specify pre installation requirements [electrical, HVAC, Compressed air ( please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour) ; if any of the above is essential]
20 | Please specify footprint size & weight.
21 | Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
22 | Predispatch inspection at factory will be required if desired by authorities.
23 | Please visit the site to check whether quoted model can be accomodated in the available space.
24 | Please specify power consumption.
25 | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
26 | Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
1 | Warranty : 2 years
2 | CMC : Maximum 5% per annum for 8 years on ex-works cost after warranty
3 | Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted.
1. If 2 years warranty, CMC maximum 5% per annum for 8 years on Ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3. Only Yes or No should be mentioned in compliance report and additional information in remarks column.