## Sheet1
Serial # 4
TMH/MUL/2021-22/CAP/PT/0057
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Dept Name:Surgery, HBCHRC, Mullanpur
Technical specifications - "Hopkin Scope 90 Degree with Light source "QTY: 6 Nos
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Year of Probable end of support:
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
A | 90 degree telescope
1 | It should be a rigid laryngo-pharyngoscope provided with handle
2 | Its diameter should be 10mm & length should be atleast 15cm
3 | Sterilizable through steam, ETO & Flash autoclavable
4 | It should avail fibre optic light transmission
B | Light Source
1 | Should have LED lamp hours of atleast 30000Hrs
2 | Should have colour temperature of 6500 degree K
3 | Light intensity should be adjustable manually
4 | Power supply- 110-240VAC, 50-60Hz
C | Light Cable
1 | Should be fiber optic with straight connector and heat resistant
2 | Diameter of not less than 4mm and length of approx 230cm
3 | Should have safety lock while connecting to the light source
Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABH norms)
Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
Please specify pre installation requirements (electrical, HVAC etc.)
Please specify footprint size & weight.
Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
Predispatch inspection at factory will be required if desired by authorities.
Please visit the site to check whether quoted model can be accomodated in the available space.
Please specify power consumption.
International standards : The unit should be CE/ FDA/BIS/ ICMED approved.
Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
Warranty: 2 years
AMC: AMC @ maximum 2% per annum for 8 years on Ex-works cost after warranty
Back to back assurance to be taken by the supplier from the OEM to supply spares for minimum 10 years and to be submitted.
1) If 2 years warranty,AMC @ maximum 2% per annum for 8 years on Ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject offer.
2) Please mention whether UPS / Stabilizer is required for the equipment. If so, should be included in the offer.
3) Only yes or no should be mentioned in the compliance report and additional info in remarks column.