## Sheet1
Serial # 4
TMH/MUL/2021-22/CAP/PT/0058
TATA MEMORIAL HOSPITAL
PURCHASE DEPARTMENT
Department - Surgery, HBCHRC Mullanpur.
Technical specification of "LED OT Light for Minor OT", Quantity 04 nos.
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Year of Probable end of support:
1 | Should have single Colour high performance LEDs with life time more than 40,000 hours.
2 | Should have minimum dual dome and the main light and satellite should have the following specifications:
a | LUX intensity 1,00,000 Lux or above & Satellite 1,20,000 Lux or above.
b | Light Field diameter should be adjustable 20-25cm
c | Colour temperature should be between 4000 to 4500 degree K
d | Colour rendering index should not be less than 95.
e | Depth of illumination should be 85 cm. or more
f | Illumination adjustment 30% to 100%
g | The light dome shall be compatible for laminar air flow
h | The light should be focusable as per the surgeon convenience
3 | Should have stable illumination throughout the life period of the light. If the intensity reduces during the warranty or CMC period the LEDs has to be replaced at free of cost if required.
4 | The LED’s must be of a single color suitable for long term maintenance and ease of replacement
5 | Temperature rise at the surgeon head level should be less than 2 degree C. Irradiance of the light should be <500 m.watts / square meter . lux as per IEC standard.
6 | Should have control panel for light focusing , intensity adjustment & ON/OFF fixed on the dome or arms and also on the walls.
7 | Should supply autoclavable handle 3 Nos. for each dome.
8 | Light should be suitable for cavity surgeries as well as superficial surgeries.
9 | Should have safety certificate from a competent authority CE or IEC60601-2. Certificates shall be produced along with the technical bid.
10 | The intensity of light from each dome should be uniform during the surgery.
11 | Unit should function with 200-240Vac, 50Hz input power supply.
12 | 360 degree rotation should be possible for main arm. And 330degree for domes.
13 | Standard accessories (Which comes along with the machine, pricing of which is part of the equipment) Please list down the Standard accessories
14 | Essential Accessories (Accessories which are required for the equipments functional needs but is not part of the standard equipment). Please list down the Essential accessories and with prices in the Financial bid.
15 | Optional accessories (List out all the optional attachments & accessories compatible with the specified model and with prices in the Financial bid)
16 | Any other salient feature you want to convey about the equipment
17 | Essential spare parts & consumables:
Provide List of required spares and consumables (if any) with their list price and % discount offered in the Financial bid
18 | Installation:
Company trained personnel as per end user’s satisfaction should do installation and impart training to end user.
19 | Training of staff:
Provide necessary training to Institute’s staff / engineers in routine maintenance and fault finding.
20 | Service support:
Complete operational manuals and technical information Company should provide technical support for hardware and software, spares and consumables.Company should provide required spares and consumables for at least 7 years even after discontinuation of the model.
21 | Safety requirements:
It should follow International safety requirement.
22 | User’s list:
A list of installations with the address and contact numbers to be provided. (User list should be for the quoted model)
Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABH norms)
Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
Please specify pre installation requirements (electrical, HVAC etc.)
Please specify footprint size & weight.
Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
Predispatch inspection at factory will be required if desired by authorities.
Please visit the site to check whether quoted model can be accomodated in the available space.
Please specify power consumption.
The unit should be CE/FDA/BIS/ICMED approved.
Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
Response Time: Should not be more than 12hrs from lodging a breakdown complaint on toll free or by email.
Warranty : 2 years
AMC : Maximum 2% per annum for 8 years on ex-works cost after warranty
Back to back assurance to be taken by the supplier from the OEM to supply spares for minimum 10 years and to be submitted.
1. If 2 years warranty, AMC maximum 2% per annum for 8 years on ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3. Only Yes or No should be mentioned in compliance report and additional information in remarks column.