## Sheet1
Serial # 4
TMH/MUL/2021-22/CAP/PT/0062
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Dept Name:Surgery, HBCHRC, Mullanpur
Technical specifications - Argon Plasma Cautery - QTY: 1 No
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Year of Probable end of support:
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
1 | The unit should produce Argon Plasma Beam for medical argon plasma coagulation
2 | The unit should be able produce different APC output when combined with electrosurgery unit.
3 | The Should have one Plug & Play Argon Plasma Coagulation socket and possibility to add another socket should be there.
4 | The unit should be available with Four different filter integrated disposable flexible endoscopic applicators for different purposes-Straight firing, Conical firing, Circumferential firing tip and double balloon endoscopy applicator.
5 | The unit should be supplied with reusable APC handle with hand switch , Open surgery APC applicator and Laproscopic APC applicator for superficial coagulation and large surface coagulation
6 | Unit should have Internal Cooling system(No exhaust fan )
7 | The unit should give four different types of APC output with Possibility to modulate Flow, voltage , time delay between pulsed output and Power settings
8 | The Should be Updatable and Upgradable unit
9 | Type: Protection Class CF, Class 1, llb
10 | All Applicators should be Plug and Play
Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABH norms)
Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
Please specify pre installation requirements (electrical, HVAC etc.)
Please specify footprint size & weight.
Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
Predispatch inspection at factory will be required if desired by authorities.
Please visit the site to check whether quoted model can be accomodated in the available space.
Please specify power consumption.
The unit should be CE/FDA/ICMED/BIS approved
Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
Warranty: 2 years
CMC: CMC @ maximum 5% per annum will be considered for all bidders for 8 years after warranty
Back to back assurance to be taken by the supplier from the OEM to supply spares for minimum 10 years and to be submitted
1) If 2 years warranty,CMC @ maximum 5% per annum for 8 years on Ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject offer.
2) Please mention whether UPS / Stabilizer is required for the equipment. If so, should be included in the offer.
3) Only yes or no should be mentioned in the compliance report and additional info in remarks column.