## Sheet1
Serial # 4
TMH/MUL/2021-22/CAP/PT/0064
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Dept Name:Surgery, HBCHRC, Mullanpur
Technical specifications - " RF based vessel sealer'' Qty: 2 Nos
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Year of Probable end of support:
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
1 | System should have a generator with footswitch for RF energy Output.
2 | System should have a touch screen display for fast setup operation and on screen diagnostics.
3 | System should be equipped with advanced temperature controlled RF energy technology that can simultaneously seal and transect vessels up to and including 7mm.
4 | System should have both Footswitch and Hand switch options.
5 | System should have a standby mode to ensure safety.
6 | System should be equipped with system diagnostics and troubleshooting guide to pin point any problems in the systems.
7 | System should have onscreen warning display system for errors.
8 | System should be able to power RF energy hand instruments for both open & laparoscopic procedures
9 | System should have the ability for software upgrade.
10 | System should comprise of the following
A | Hardware:
1 Generator
2 Footswitch & Cable
B | Accessories:
Instruments for Open & Lap (Compatible with quoted model)
C | Instruments:
List of Instruments ( RF energy instruments for open & Lap) with price to be provided & it should be compatible with quoted model.
Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABH norms)
Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
Please specify pre installation requirements (electrical, HVAC etc.)
Please specify footprint size & weight.
Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
Predispatch inspection at factory will be required if desired by authorities.
Please visit the site to check whether quoted model can be accomodated in the available spece.
Please specify power consumption.
International standards : The unit should be CE/ FDA/BIS/ICMED approved
Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
Warranty: 2 years
CMC: CMC @ maximum 5% per annum for 8 years on Ex-works cost after warranty
Back to back assurance to be taken by the supplier from the OEM to supply spares for minimum 10 years and to be submitted to TMH .
1) If 2 years warranty,CMC @ maximum 5% per annum for 8 years on Ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject offer.
2) Please mention whether UPS / Stabilizer is required for the equipment. If so, should be included in the offer.
3) Only yes or no should be mentioned in the compliance report and additional info in remarks column.