## Neuro
Serial # 4
TMH/MUL/2021-22/CAP/PT/0065
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Dept Name:Surgery, HBCHRC, Mullanpur
Technical specifications - "Electric High speed drill" OTY : 02 Nos
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Yes | No | Remark
Model Name:
Make:
Year of Introduction:
Year of Probable end of support:
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
1 | Electric Console
a | Should have console for drill handpiece system.
b | Should have function of controlling brightness, contrast and alarm volumes on the console.
d | Console should offer change from variable/fixed mode, forward/reverse.
e | Console should offer change of speed.
f | Console should offer change in irrigation percentage.
g | Integrated irrigation pump should be available for drill .
2 | Footswitch
It should have:
a | Should have fully programmable footswitch as user needs
b | User should be able to control following functions via footswitch:
> Forward,reverse, oscillation
> Switch over to high/low speed
> Increase/decrease irrigation flow rate
3 | High speed micro drill motor:
It should have:
a | Maintenance free, brushless motor
b | Handpiece speed should be atleast 70,000 RPM or more
c | Sterilizable through steam, ETO, flash autoclavable.
d | Direction of rotation should be both forward & reverse.
4 | Attachments required:
a | Neuro Craniotomy attachments
b | Endo Neuro attachments
5 | Connecting Cord
1 | Should beatleast 10ft long, 3/8'' diameter flexible electrical connecting cord.
2 | Should be autoclavable
Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABH norms)
Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
Please specify pre installation requirements (electrical, HVAC etc.)
Please specify footprint size & weight.
Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
Predispatch inspection at factory will be required if desired by authorities.
Please visit the site to check whether quoted model can be accomodated in the available space.
Please specify power consumption.
The unit should be CE /FDA/ BIS/ ICMED approved
Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
Warranty: 2 years
CMC: CMC @ maximum 5% per annum for 8 years on Ex-works cost after warranty
Back to back assurance to be taken by the supplier from the OEM to supply spares for minimum 10 years and to be submitted
1) If 2 years warranty,CMC @ maximum 5% per annum for 8 years on Ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject offer.
2) Please mention whether UPS / Stabilizer is required for the equipment. If so, should be included in the offer.
3) Only yes or no should be mentioned in the compliance report and additional info in remarks column.