## Sheet2
Serial # 4
TMH/TMC/2021-22/CAP/PT/0066
TATA MEMORIAL HOSPITAL
PURCHASE DEPARTMENT
Department - CSSD, HBCHRC-Mullanpur, HBCHRC-Vizag and MPMMCC-Varanasi.
Technical Specifications of "Low Temperature based H2O2 Gas Sterilizer / Plasma Sterilizer ", Quantity 03 nos. (Mullanpur Qty: 01 no, Vizag Qty : 01 no, MPMMCC Varanasi Qty: 01 no)
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Name of the Model:
Make:
Specification:
Year of Introduction:
Probable end of support:
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
1 | Simple & fast low temperature sterilization with the help of H2O2 & effective removal of H2O2 from sterilized items and to compliment the process
2 | Should be suitable for sterilization of medical items like rigid endoscopes, lumen & non lumen , metal , non metal, heat & moisture sensitive instruments
3 | Chamber should have usable volume of 90 litres or above with rectangular chamber with usable space of 1 STU
4 | The sterilization temperature inside the chamber should be less than 55°C
5 | Cycle time should be 35 to 60 mins
6 | The sterilant should be in a cassette/ bottle with H2O2 concentration 55% ± 10%
7 | System should be recommended by Intuitive surgical . Recommendation letter should be submitted along with the bid. Should be able to effectively sterilize MIS instruments, scopes, etc.
8 | The quoted model should be indicated for sterilization of metal and non-metal medical devices. This compliance should be certified by CE/ FDA/ICMED/ BIS.
9 | Lumen sterilization claims should be validated & endorsed. This compliance should be certified by CE/FDA/BIS/ ICMED
10 | The moisture detection phase/ load conditioning phase should be available in the equipment prior to sterilant injection phase to reduce aborts due to moisture and avoids sterilant wastage. Valid documents to be provided for the same
11 | The system should use minimum quantity of sterilant which should be less than 6-8 ml per injection to deliver dry terminal sterilization to ensure safety of Instruments against corrosion.
12 | The sterilizer should have ergonomic foot pad for convenience and safety while carrying heavy loads.
13 | Sterilant (H2O2 concentration >= 55%) should be in cassette/bottle/cup . Should have safety mechanism to ensure sterilant leakages. Please specify.
14 | Sterilant cassette/bottle/Cup should be able to store at room temperature for atleast 10 months from the date of receipt at TMC i.e same shelf life of 10 months applicable to be in usable condition at TMC will be considered otherwise it should be replaced at No cost to TMC
15 | Should be environment friendly and by-products should be non toxic.
16 | Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
17 | Please specify pre installation requirements [electrical, HVAC, Compressed air ( please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour) ; if any of the above is essential]
18 | Please specify footprint size & weight.
19 | General Terms:
a | Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
b | Predispatch inspection at factory will be required if desired by authorities.
c | Please visit the site to check whether quoted model can be accomodated in the available space.
d | Please specify power consumption.
e | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
f | Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
g | Uptime Guarantee- 95 %
h | Response time: call need to be attended within 12 hours from call log or email.
i | Penalty clause: The supplier and/or its Indian agent will be required to give the guarantee towards the performance of the equipment during the warranty period and in case of comprehensive annual maintenance contract period for maintaining the equipment in good working condition for a period of at least 300 days out of a period of 365 days a year (i.e. 95% uptime) 24X7. The ten hours non-functioning of the equipment or any part thereof will be considered as one day down time. Total 3000 hours in a year will be considered as Uptime guarantee. Working time is 8.00 a.m. to 6.00 p.m. on week days and 8.00 a.m. to 4.00 p.m. for Saturdays. The equipment shall be fully functional as part of the uptime and if it is partly functional it shall be considered as downtime and deduction shall be made on Pro-rata basis i.e. if the equipment is 70% functional 30% downtime shall be applicable. The decision of the TMC management or its representative in determining the % of the downtime shall be final and binding. Essential period to shut down the installation entirely or partially shall also be included in the down time while calculating the 95% guaranteed uptime i.e. all features as per specifications in purchase order should be functional for uptime. The supplier / agent shall be required to pay a penalty for every day 0.1% of total cost of the equipment per day per machine excluding 150 hours permissible downtime. This rate shall increase commensurate with the rate increase of CMC. If the number of days downtime in each period of 365 days is more than the downtime permissible (which is to be calculated as defined above), the delay for not bringing the equipment in functioning order in any way, directly or even partially will be attributable to the firm. The bidder should provide the details of preventive maintenance to be undertaken through the year. Preventive maintenance to be done on holidays only. Preventive maintenance and breakdown will not be clubbed. The bills should be raised quarterly and will be paid at the end of the quarter only.
j | Warranty:2 years
k | CMC - 5% of ex works cost per annum will be considered for all bidders for 8 years. (Please specify the list & cost of spare parts , consumables/ consumable spares (i.e spares need to be replaced at regular intervals, may be quarterly/half yearly/ yearly such as annual maintenance kit etc.) if any. All such items should be included in CMC.
l | Life cycle cost of the equipment and per cycle consumable cost will be considered to arrive at L1
m | Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted to
1. If 2 years warranty, 8 years CMC and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3. Only Yes or No should be mentioned in compliance report and additional information should in remarks column.