## Reviced Specs
Serial # 4
TMH/MUL/2021-22/CAP/PT/0067
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
DEPARTMENT: Surgery, HBCHRC Mullanpur
Technical Specifications of "Cystoscopy cart (Urology)"QTY-01 No
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr no | Description | Qty | Compliance |  | Remarks
|  |  | Yes | No
Model Name:
Make:
Year of Introduction:
Probable end of support:
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
1 | Full high definition(HD) camera system | 1pc
A. Camera control unit
1. Outputs- 2 x DVI 1080p ,2 x HD SDI 1080
2. DVI , HD ( length 3m)should be supplied
3. power supply- 230VAC, 50-60Hz
B. Camera head with coupler
1. Image sensor Should be of 1 x 1/3'' or 3 x 1/3'' & CCD / CMOS with atleast 2 head camera buttons for zoom, focus, white balance etc.
2. Image resolution- 1920 x 1080p
3. Endoscope coupler for standard eyepiece should be detachable, focal length Upto 30mm
4. Camera cable should be atleast 3m long
2 | LED Light Source | 1pc
a. High performance LED
b. LED lamp life should be 25000hrs
c. Power supply- 230V ac/ 50-60Hz
d. Should have lamp life over indicator
3 | Fibre Optic Cable | 2pc
a. Should be heat resistant, have enhanced light transmission
b. Diameter 4.8mm, length 230cm
c. Should be autoclavable & have telesope adaptor & light adaptor
4 | High Definition (HD) LED monitor | 1pcs
a. Display Technology- Full HD
b. Image resolution- 1920 X 1080
c. Screen size- 26/27inches
d. Image Aspect ratio- 16:9
5 | Recording Unit
1 | Should be a modular and preconfigured system for intraoperative documentation of still images and videos.
2 | Should record still images and videos in full HD ( with resolution 1920 X 1080) with no degradation in image quality.
3 | Video recording format should be MPEG 4 AVC/H.264
4 | Still image recording format should be BMP/TIFF/JPEG
5 | Compatible media should be Internal HDD/Blu-ray disks/DVD/USB stick
6 | Internal HDD capacity should be atleast 320GB
7 | Should have structured & clear user guidance to fill patient's data (ID, procedure name, sex, Surgeon's name , etc), and have access for functions like play, recording start-stop, pause, capture.
8 | Supported color system should be PAL, NTSC
9 | Should have an in-built display monitor or separate HD monitor should be provided which allows to record or capture in real time
10 | Should have video input signals such as DVI, Y/C, Composite
11 | Serial remote- RS 232C, stereo mini pin jack
12 | Power Supply- 230VAC, 50-60Hz
6 | Trolley with electrical point & multiple shelves | 1pc
7 | Telescope 4mm, 30° | 3 pcs
a. Should be autoclavable, have enlarged view
b. Diameter should be 4mm, length 30cm.
c. fibre optic light transmission should be incorporated
8 | Telescope 4mm, 0° | 1pc
a. Should be autoclavable, have enlarged view
b. Diameter should be 4mm, length 30cm
c. fibre optic light transmission should ne incorporated
9 | Cystoscopy Sheath with obturator 20Fr/21Fr | 1 pcs
10 | Cystoscopy Sheath with obturator 17Fr | 2 pcs
11 | Cystoscopy Sheath with obturator 22Fr/23Fr | 1 pcs
12 | Resectoscopy sheath 26 Fr | 2pcs
a. Should have connecting tubes for in & outflow
b. Should have oblique beak with standard obturator
c. Rotatable inner tube should be provided
13 | Working Element
| Cutting should be by means of finger grip. The thumb support is not movable. In rest position, the electrode is outside the sheath | 1 pcs
| Cutting should be by means of thumb grip. The finger support is not movable. In rest position, the electrode is outside the sheath | 1 pcs
Working element set should consist of:
| a. Cutting loops | 10pcs
| b. Coagulating electrodes | 5pcs
| c. Collins Knife | 5pcs
| d. High frequency Cords | 4pcs
| e. VIU knife | 4pcs
| f. Halfround knife | 1pc
| g. Visual Obturator | 1pc
14 | Optical urethrotomy Sheath with half round sheath | 1pc
a. Obturator should be supplied
15 | Optical Biopsy Forceps Rigid- should have double action jaw | 2pc
16 | Cystoscopy flexible Biopsy Forceps 7 Fr (Oval Spoon) | 2pc
a. Should be flexible, double action jaw, 7 Fr, length 40cm
17 | Telescope bridge with 1 instrument channel | 3pcs
18 | Telescope bridge with 2 instrument channel | 1pc
19 | Cystoscopy Grasping Forceps 7 Fr- with Alligator jaw, length 40cm, Flexible-2 , Rigid-1 | 3pcs
20 | Urovac bladder evacuator | 2pcs
21 | Toomey Syringe (50cc) | 2pcs
22 | Urethral dilator set male & female | 1pc
23 | Cystoscopy penile clamp | 5pcs
24 | OTIS Urethrotome | 1pc
25 | OTIS Urethrotome blade | 2pcs
26 | All the instruments should be CE or FDA or BIS or ICMED Certified and of Single parent company except Trolley , Recording Device, HD Monitor, Urovac,toomey syringe,urethral dilator set and cystoscopy penile clamp.
The unit should be CE /FDA/ ICMED/ BIS approved
Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, cosumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABL norms)
Please specify list of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
Please specify pre installation requirements ( electrical, HVAC etc.)
Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
UPS of adequate capacity to be provided if required.
Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
Warranty: 2 years
CMC- CMC @ maximum 5% per annum for 8 years on Ex-works cost after warranty
Back to back assurance to be taken by the supplier from the OEM to supply spares for minimum 10 years and to be submitted to TMH.
1) | If 2 years warranty and CMC @ maximum 5% per annum for 8 years on Ex-works cost after warranty conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2) | Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3) | Only Yes or No should be mentioned in compliance report and additional information should in remarks column.