## final
Serial # 4
TMH/MUL/2021-22/CAP/PT/0068
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Dept Name: Surgery, HBCHRC, Mullanpur
Technical specifications - Video transnasal Oesophagoscope - QTY: 1No
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr No | Item Description | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Probable end of support:
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
A) | Video transnasal Oesophagoscope
1 | Should have the facility of simultaneous suction and application of medication via removable suction valve.
2 | It should be high definition, waterproof, fully immersible for cleaning and disinfection and Sterilizable.
3 | Tip deflection should be 210 UP/90 Down deg or more.
Tip deflection should be Right/Left 100 deg or more.
4 | Angle of view should be more than or equal to 140 degrees.
5 | Working length should be more than or equal to 1100 mm.
6 | Working channel diameter should be of size from 2.0 to 2.5mm.
7 | Outer diameter of the distal end should be of size from 5.0mm to 6.0 mm.
8 | Insertion tube outer diameter should be of size from 5.0mm to 6.0mm.
9 | Depth of field should be of size from 4-100mm or better.
10 | It should include the standard accessories as Carrying case, leakage tester, bite protector, cleaning brush,Biopsy Forceps.
11 | Two or more no. of remote control switches on control body.
12 | Scope should be fiber less with CCD on tip , compatible with video processor.
13 | Should be able to ensuffulate air .
B) | Specs for video Processor
1 | Compatible with Analog and Digital (Y/C,RGB, COMPOSITE,DVI) to reproduce high definition (1080p/i) images / video.
2 | White balance adjustment should be available
3 | Colour adjustment (Red, Green, Blue ) should be available
4 | It should have facility for Image Enhancement Endoscope iscan/NBI OR FICE technology for better diagnosis.
C) | Specs for Light Source
1 | Lamp type should be LED(Equivalent to Xenon)
2 | Average lamp life should be 10000 Hrs or more
3 | power supply- 230VAC, 50Hz
4 | Type of protection against electric shock- Class I
5 | It should have provision to adjust the light intensity as per needs
6 | It should be provided with fibre optic cable of atleast 3m long or more
D) | Atleast 26” medical grade Monitor to be provided with complete system for visualisation.
E) | Image management and recording software (PC based or built in) to supplied, along with necessary hardware & atleast 1TB storage
F) | Should be supplied with cart/trolley. (Indian Make)
1 | International standards : The unit should be CE/ FDA/BIS/ ICMED approved.
2 | Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, cosumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABL norms)
3 | Please specify list of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
4 | Please specify pre installation requirements ( electrical, HVAC etc.)
5 | Please specify the footprint size & weight
6 | Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
7 | Predispatch inspection at factory will be required if desired by authorities
8 | Please visit the site to check whether quoted model can be accomodated in the available space.
9 | Please specify power consumption
10 | All necessary standards and norms. (As applicable)
11 | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls. Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
12 | Warranty: 2 years
13 | CMC- CMC @ maximum 5% per annum for 8 years on Ex-works cost after warranty
14 | Back to back assurance to be taken by the supplier from the OEM to supply spares for minimum 10 years and to be submitted to HBCHRC Mullanpur
1) | If 2 years warranty and CMC @ maximum 5% per annum for 8 years on Ex-works cost after warranty conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2) | Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3) | Only Yes or No should be mentioned in compliance report and additional information should in remarks column.