## Mullanpur
Serial # 4
TMH/TMC/2021-22/CAP/PT/0069
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
DEPARTMENT: Surgery, HBCHRC Mullanpur
Technical Specifications of "Endoscopy System "QTY: 01 No
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Probable end of support:
A | Equipment Cart:
1 | It should be wide enough and high to occupy monitor, light source, camera control unit
2 | It should have 4 dual antistatic wheels, 2 equipped with locking brakes
3 | It should have integrated electrical subdistributors with minimum 6 sockets
B | Camera control Unit
1 | It is to be used with three chip HD camera head
2 | It should have integrated digital image processing module
3 | It should offer maximum resolution as 1920 X 1080 pixels
4 | It should support the colour systems like NTSC/PAL
5 | It should have all the video output such as SDI, DVI-D, etc
6 | Type of protection against electric shock- Class I
7 | Power supply- 230VAC, 50Hz
8 | Should have image enhancement technology for early detection malignancy and aid surgery.
C | Camera Head
1 | It should be a 3 chip camera head
2 | Weight Should not be more than 350 gms
3 | Focal Distance should be in the range of 16mm -31 mm
4 | It should have minimum 2 fully programmable head buttons
5 | It should support the colour systems like NTSC/ PAL
6 | It should offer full HD resolution with excellent colour reproduction
7 | It should incorporate power zoom & power focus which can be controlled with one hand.
8 | Pixel output should be 1920 X 1080
D | Light Source
1 | Lamp type should be LED (Equivalent to Xenon 300W)
2 | Average lamp life should be 10000Hrs or more.
3 | power supply- 230VAC, 50Hz
4 | Type of protection against electric shock- Class I
5 | It should have provision to adjust the light intensity as per needs
6 | It should be provided with fibre optic cable of atleast 2m long or more
F | Medical Grade Monitor
1 | It should be a 26'' full HD monitor.
2 | Resolution should be 1920 X 1080 Pixels
3 | Aspect ratio- 16:9
4 | It should support the colour systems like NTSC/ PAL
5 | It should have all types of video inuts such as DVI, SDI, HD-SDI, S-video, Composite video
6 | It should have all types of video outputs such as DVI, SDI, HD-SDI, S-video, Composite video
7 | Monitor screen must be with anti reflection coated glass.
8 | It should include the following:
a. Power supply (230 VAC, 50Hz)
b. Mains cord
c. 1 connecting cable of each input type at least 3meters long
d. Monitor stand
9 | Please mention make & model name of the monitor
G | Recording Unit
1 | Should be a modular and preconfigured system for intraoperative documentation of still images and videos.
2 | Should record still images and videos in full HD ( with resolution 1920 X 1080) with no degradation in image quality.
3 | Video recording format should be MPEG 4 AVC/H.264
4 | Still image recording format should be BMP/TIFF/JPEG
5 | Compatible media should be Internal HDD/Blu-ray disks/DVD/USB stick
6 | Internal HDD capacity should be atleast 500GB
7 | Should have basic functions for editing of still images & videos
8 | Should have integrated file viewer for still images & videos
9 | Should have structured & clear user guidance to fill patient's data (ID, procedure name, sex, Surgeon's name , etc), and have access for functions like play, recording start-stop, pause, capture.
10 | Supported color system should be PAL, NTSC
11 | Should have an in-built display monitor or separate HD monitor should be provided which allows to record or capture in real time
12 | Should have video input signals such as DVI, Y/C, Composite
13 | Serial remote- RS 232C, stereo mini pin jack
14 | Power Supply- 230VAC, 50-60Hz
H | Slave monitor
1 | It should be a 26" full HD medical grade monitor
2 | Resolution should be 1920 x 1080 pixels
3 | Aspect ratio- 16:9
4 | It should support the colour systems like NTSC/PAL
5 | It should have all type of video inputs such as DVI, SDI, HD-SDI, S- video composite video
6 | It should have all type of video outputs such as DVI, SDI, HD-SDI, S- video composite video
7 | Monitor scaeen must be with anti reflection coated glass.
8 | It should include the following:
a. Power Supply (230 VAC, 50Hz)
b. Mains Cord
c. 1 connecting cable of each input type at least 3 meters long
d. monitor Stand
9 | Please mention make & model name of the monitor
Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, cosumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABL norms)
Please specify list of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
Please specify pre installation requirements ( electrical, HVAC etc.)
Please specify foot print size & weight.
Demo of the quoted model, will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
Predispatch inspection at factory will be required if desired by authorities.
Please visit the site to check whether quoted model can be accomodated in the available space.
Please specify power consumption.
UPS of adequate capacity to be provided if required.
The unit should be CE/FDA/BIS/ICMED approved.
Back to back assurance to be taken by the supplier from the OEM to supply spares for minimum 10 years & to be submitted.
Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
Response Time: Should not be more than 12hrs from lodging a breakdown complaint on toll free or by email.
Warranty: 2 years
CMC: 5% of ex workss cost per annum to be consided for all bidders for 8 years after warranty
1) | If 2 years warranty, CMC 5% of ex workss cost per annum to be consided for all bidders for 8 years after warrantyand buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2) | Please mention whether UPS/ stabiliser is required for the equipment. If so should be in the offer
3) | Only Yes or No should be mentioned in compliance report and additional information should in remark columns.

## Vizag
Serial # 4
TMH/TMC/2021-22/CAP/PT/0069
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
DEPARTMENT: Surgery, HBCHRC Vizag
Technical Specifications of "Endoscopy System "QTY: 01 No
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Probable end of support:
A | Equipment Cart:
1 | It should be wide enough and high to occupy monitor, light source, camera control unit
2 | It should have 4 dual antistatic wheels, 2 equipped with locking brakes
3 | It should have integrated electrical subdistributors with minimum 6 sockets
B | Camera control Unit
1 | It is to be used with three chip HD camera head
2 | It should have integrated digital image processing module
3 | It should offer maximum resolution as 1920 X 1080 pixels
4 | It should support the colour systems like NTSC/PAL
5 | It should have all the video output such as SDI, DVI-D, etc
6 | Type of protection against electric shock- Class I
7 | Power supply- 230VAC, 50Hz
8 | Should have image enhancement technology for early detection malignancy and aid surgery.
C | Camera Head
1 | It should be a 3 chip camera head
2 | Weight Should not be more than 350 gms
3 | Focal Distance should be in the range of 16mm -31 mm
4 | It should have minimum 2 fully programmable head buttons
5 | It should support the colour systems like NTSC/ PAL
6 | It should offer full HD resolution with excellent colour reproduction
7 | It should incorporate power zoom & power focus which can be controlled with one hand.
8 | Pixel output should be 1920 X 1080
D | Light Source
1 | Lamp type should be LED (Equivalent to Xenon 300W)
2 | Average lamp life should be 10000Hrs or more.
3 | power supply- 230VAC, 50Hz
4 | Type of protection against electric shock- Class I
5 | It should have provision to adjust the light intensity as per needs
6 | It should be provided with fibre optic cable of atleast 2m long or more
F | Medical Grade Monitor
1 | It should be a 26'' full HD monitor.
2 | Resolution should be 1920 X 1080 Pixels
3 | Aspect ratio- 16:9
4 | It should support the colour systems like NTSC/ PAL
5 | It should have all types of video inuts such as DVI, SDI, HD-SDI, S-video, Composite video
6 | It should have all types of video outputs such as DVI, SDI, HD-SDI, S-video, Composite video
7 | Monitor screen must be with anti reflection coated glass.
8 | It should include the following:
a. Power supply (230 VAC, 50Hz)
b. Mains cord
c. 1 connecting cable of each input type at least 3meters long
d. Monitor stand
9 | Please mention make & model name of the monitor
G | Recording Unit
1 | Should be a modular and preconfigured system for intraoperative documentation of still images and videos.
2 | Should record still images and videos in full HD ( with resolution 1920 X 1080) with no degradation in image quality.
3 | Video recording format should be MPEG 4 AVC/H.264
4 | Still image recording format should be BMP/TIFF/JPEG
5 | Compatible media should be Internal HDD/Blu-ray disks/DVD/USB stick
6 | Internal HDD capacity should be atleast 500GB
7 | Should have basic functions for editing of still images & videos
8 | Should have integrated file viewer for still images & videos
9 | Should have structured & clear user guidance to fill patient's data (ID, procedure name, sex, Surgeon's name , etc), and have access for functions like play, recording start-stop, pause, capture.
10 | Supported color system should be PAL, NTSC
11 | Should have an in-built display monitor or separate HD monitor should be provided which allows to record or capture in real time
12 | Should have video input signals such as DVI, Y/C, Composite
13 | Serial remote- RS 232C, stereo mini pin jack
14 | Power Supply- 230VAC, 50-60Hz
Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, cosumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABL norms)
Please specify list of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
Please specify pre installation requirements ( electrical, HVAC etc.)
Please specify foot print size & weight.
Demo of the quoted model, will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
Predispatch inspection at factory will be required if desired by authorities.
Please visit the site to check whether quoted model can be accomodated in the available space.
Please specify power consumption.
UPS of adequate capacity to be provided if required.
The unit should be CE/FDA/BIS/ICMED approved.
Back to back assurance to be taken by the supplier from the OEM to supply spares for minimum 10 years & to be submitted.
Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
Response Time: Should not be more than 12hrs from lodging a breakdown complaint on toll free or by email.
Warranty: 2 years
CMC: 5% of ex workss cost per annum to be consided for all bidders for 8 years after warranty
1) | If 2 years warranty, CMC 5% of ex workss cost per annum to be consided for all bidders for 8 years after warrantyand buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2) | Please mention whether UPS/ stabiliser is required for the equipment. If so should be in the offer
3) | Only Yes or No should be mentioned in compliance report and additional information should in remark columns.