## Sheet1
TMH/TMH/2022-23/CAP/LT/0031
TATA MEMORIAL HOSPITAL
PURCHASE DEPARTMENT
Department - Dental & Prosthetics Surgery , TMH.
Technical specification of "Radio Visuo Gram (RVG)", Quantity 01 no.
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description | Complaince |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Probable end of support:
1 | Should protect from over exposure and ensures perfect and clear images in first shot
2 | Should be possible to take x-ray in single click
3 | Should be of CMOS generation
4 | The Fibre optic should be able to protect the sensor against X-rays and increase its durability
5 | Should have automatic control of exposure
6 | Should have easy positioning of sensor inside the mouth/ thickness should be less than or equal to 5mm
7 | Should also be easier to place the X-ray cone perpendicular to the x-ray receptor
8 | Should have 25 lp/mm Theoretical resolution.
9 | Should have a pixel size of 20 μm x 20 μm.
10 | Should have 3 to 4 m cable length for the USB cable. Extension should be provided otherwise.
11 | Should be provided with Imaging Software (Free software support including reinstallations as & when required till RVG hardware end of life support by OEM)
12 | Should have Standard Size 1 sensor (equivalent to adult IOPA film or similar)
13 | Should have the active surface area of 600 mm² (20 x 30 mm)
14 | Should be supplied with sensor holder & disposable sensor covers
15 | Should be supplied with latest PC configuration i5 or better , 3.7GHz 6C processor (12MB Smart 8 GBDDR4 (2400MHz or higher)/4 DIMM Slots/1TB/No ODD/ Win 10 Pro/85% efficiency / 3 video ports (VGA + 2DP), expandable /serial port/4 USB 3.1 port & 4 USB2.0 ports/ 19.5 inches LED colour monitor
16 | Should be CE/FDA/BIS/ICMED approved
17 | Should perform yearly calibration/Validation and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABL/NABH norms). To submit SOPs for PMS at the time of installation/commissioning.
18 | Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
19 | Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
20 | Please specify pre installation requirements [electrical, HVAC, Compressed air ( please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour) ; if any of the above is essential]
21 | Please specify footprint size & weight.
22 | Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
23 | Predispatch inspection at factory will be required if desired by authorities.
24 | Please visit the site to check whether quoted model can be accomodated in the available space.
25 | Please specify power consumption.
26 | Please quote cost for buy back for Trophy C1060 ( Make :- Trophy ) on "as is where is basis", Qty - 01 no.
27 | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
28 | Response Time: Should not be more than 06 hrs from logging a breakdown complaint on toll free or by email.
1 | Warranty- 2 years
2 | AMC- 2% of Ex- Works cost per annum will be considered for all bidders for 8 years after warranty
3 | Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted.
1. If 2 years warranty, AMC maximum 2% per annum for 8 years on ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3. Only Yes or No should be mentioned in compliance report and additional information in remarks column.