## final
TMH/VIZ/2022-23/CAP/LT/0034
TATA MEMORIAL HOSPITAL
PURCHASE DEPARTMENT
Department - Nursing, HBCHRC Vizag
Technical specification of "Air Mattress", Quantity 10 nos.
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr. No. | Description of Item | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Year of Probable end of support:
1 | There should be controls, alarms and indicators.
2 | Regular inflation deflation of aircells redistributing pressure automatically.
3 | Constant lower pressure on vulnerable areas with pressure sensitivity.
4 | Preferable if adjustable pressure as per weight range of patient.
5 | Should be useful for prevention as well as pressure ulcer treatment protocols.
6 | Detachable cover if needs to be disinfected.
7 | CPR control mode should be there.
8 | Mode of Operation: Continuous or Static .
9 | Maximum patients weight capacity of 180-200 kg.
10 | Should be able to treat patient with Medium to high Risk or up to Grade 3 pressure ulcer.
11 | Should be able to transport along with patient mattress without pump with simple operation via disconnection or tube-set from the pump.
12 | Should be provided with standard mattress and the connecting tube between the pump and the mattress.
13 | Power supply 100-240 VAC , 50 -60 HZ.
14 | Standards : The unit should comply with standards and should have CE or FDA marking / BIS / ICMED approved.
15 | Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABL norms). To submit SOPs for PMS at the time of installation/commissioning.
16 | Performance certificates of IQ,OQ,PQ tests to be submitted at the time of installation/commissioning ; if desired by the end user. (Optional)
17 | Should be supplied with standard make UPS with atleast 30 mins backup. (Optional)
18 | Please submit the technical data sheet along with the brochure as a tender document with reference to the quoted make & model.
19 | Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
20 | Please specify pre installation requirements [electrical, HVAC, Compressed air ( please specify required air pressure), water requirement (RO/DI/Distilled water with its pressure, flow rate per hour) ; if any of the above is essential]
21 | Please specify footprint size & weight.
22 | Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
23 | Predispatch inspection at factory will be required if desired by authorities.
24 | Please visit the site to check whether quoted model can be accomodated in the available space.
25 | Please specify power consumption.
26 | Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
27 | Response Time: Should not be more than 12 hrs from lodging a breakdown complaint on toll free or by email.
1 | Warranty- 2 years
2 | AMC- 2% of FOB cost per annum will be considered for all bidders for 8 years after warranty
3 | Back to back assurance to be taken by the supplier from OEM to supply spares for minimum 10 years and to be submitted.
1. If 2 years warranty, AMC maximum 2% per annum for 8 years on ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3. Only Yes or No should be mentioned in compliance report and additional information in remarks column.