## TMH
Serial # 4
TMH/TMC/2021-22/CAP/PT/0048
TATA MEMORIAL HOSPITAL
PURCHASE DEPARTMENT
DEPARTMENT: Pulmonary Medicine, TMH
Technical specifications - EBUS SYSTEM - QTY: 1 No
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr No | Item Description | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Probable end of support:
A) | Ultrasound Linear BRONCHOSCOPE - 01No
1 | Should have the facility of simultaneous suction and application of medication via removable suction valve.
2 | It should be waterproof, fully immersible for cleaning and disinfection and Sterilizable.
3 | Tip deflection should be 130 UP/70 Down deg or better.
4 | Angle of view should be more than or equal to 80 degrees with a forward oblique view should not be more than 45 Degree . Please specify the degree of forward view.
5 | Working length should be 60 cm or better.
6 | Working channel diameter should be of size 2.0 mm or better and it should be compatible for TBNA needle.
7 | Outer diameter of the distal end should be of size from 6.5mm to 7.0 mm.
8 | Insertion tube outer diameter should be of size from 6.0mm to 6.5mm.
9 | depth of field of 3mm- 50mm or more
10 | Ultrasound display modes should be B Mode, Color Mode.
11 | It should have scanning frequency range from 5MHz to 12 MHz or more switchable
12 | It should have scanning range 60 degree or better.
13 | Scanning direction should be Longitudinal with Scanning System – Convex
14 | It should be designated for both Balloon method & direct contact.
15 | It should include the standard accessories as Carrying case, leakage tester, bite protector, cleaning brush, balloon applicator.
16 | Two or more no. of remote control switches on control body.
17 | Scope should be compatible with video processor & Ultrasound processor.
B) | Ultrasound Radial Probe - 01No
1 | Scanning Method : Mechanical Radial
2 | It should be waterproof, fully immersible for cleaning and disinfection and Sterilizable.
3 | Scan Range : 360° circumferential, in perpendicular direction with respect to probe insertion direction
4 | It should have scanning frequency range 20MHz or better
5 | Outer diameter of the distal end should be of size from 1.4mm to 1.9 mm.
6 | It should include the standard accessories as Carrying case.
7 | Should be compatible with all 2.0mm or above bronchoscopes
8 | Should be supplied with radial probe driving unit and radial processor if applicable.
C) | DIGITAL ULTRASOUND SCANNER
1 | It should be Compact .
2 | Dual mode capabilities to see in real time both B mode and CFM mode .
3 | Elastography facility should be available to visualize tissue strain in real time, offering information of its stiffness.(Optional)
4 | High resolution imaging with HI COMPOUND TECHNOLOGY ( with adjustable modes) to transmit and receive ultrasound beams at multiple angles and overlaps the images in real time which observes clear lesions through the reduction of speckle noise and improvement of contrast resolution.
5 | Ergonomic operation keyboard for ease of operator use.
6 | It should have annotation, arrow mark and pointer display.
7 | USB Should be available for external storage facility.
8 | Tissue Harmonic Imaging that uses broadband technology for better penetration and resolution should be available for in-depth and good resolution in image quality. This technology basically expands the bandwidth of harmonic signals to maximum, making it possible to include the low frequency subharmonic signals.
9 | Should be capable of giving multiple frequency range (5 Mhz to 12Mhz) or more
10 | It should have GAIN, Contrast, STC functions
11 | Cine Memory should be 150 frames or more
D) | Specs for video Processor
1 | Compatible with Analog and Digital (Y/C,RGB, COMPOSITE,DVI,HD-SDI) to reproduce high definition (1080p/i) images / video.
2 | White balance adjustment should be available
3 | Colour adjustment (Red, Green, Blue ) should be available
4 | It should have facility for Image Enhancement Endoscope iscan/NBI OR FICE technology for better diagnosis.
E) | Specs for Light Source
1 | It should be a 300W Xenon or Multi LED light soure
2 | Lamp life should be 500hrs for Xenon and 10000 hrs for LED
F) | Medical Grade Monitor- 2Nos
1 | Atleast 26” medical grade Monitor to be provided with complete system for visualisation-2Nos separate vision for Ultrasound image and bronchosciopy image.
G) | Recording System
1 | Should provide stand alone recorder with inbuilt or separate display monitor. Should be with 1 TB internal storage and 2 TB External HDD.
2 | Should have dicom and PACS compatibility for recording system.
H) | Should be supplied with cart/trolley.
I) | Accessories
1 | Balloon: 50Nos
2 | TBNA needle (EBUS): 50Nos
3 | EBUS transbronchial biopsy needle-50Nos
J) | Optional Items
1 | Quote rate seperately for elestography
2 | Quote rate seperately for Radial probe.
3 | Quote rate seperately for Radial Biopsy Guilded sheath
4 | Quote rate seperately for Xenon and Multi LED Light source.
5 | Quote rate seperately for Multi LED lamp.
Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABH norms)
Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
Please specify power consumption.
International standards : The unit should comply with international standards and should have European CE or US FDA marking or BIS marking or ICMED
Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
Response Time: Should not be more than 6 hrs from lodging a breakdown complaint on toll free or by email.
Warranty: 2 years
CMC- CMC @ maximum 5% per annum for 8 years on Ex-works cost after warranty
Back to back assurance to be taken by the supplier from the OEM to supply spares for minimum 10 years and to be submitted.
1. If 2 years warranty, CMC @ maximum 5% per annum for 8 years on Ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3. Only Yes or No should be mentioned in compliance report and additional information should in remarks column.

## Mullanpur
Serial # 4
TMH/TMC/2021-22/CAP/PT/0048
TATA MEMORIAL HOSPITAL
PURCHASE DEPARTMENT
DEPARTMENT: Surgery, HBCHRC Mullanpur
Technical specifications - EBUS SYSTEM - QTY: 1 No
Please quote the model which is including the latest state of art equipment and meets the basic minimum specifications. The quoted model should be latest introduced with end of life not before 12 years from the date of installation at TMC.
Sr No | Item Description | Compliance |  | Remarks
|  | Yes | No
Model Name:
Make:
Year of Introduction:
Probable end of support:
A) | Adult video BRONCHOSCOPE -01No
1 | Should have the facility of simultaneous suction and application of medication via removable suction valve.
2 | It should be waterproof, fully immersible for cleaning and disinfection and Sterilizable.
3 | Tip deflection should be 130 UP/70 Down deg or better.
4 | Angle of view should be more than or equal to 120 degree.
5 | Working length should be more than or equal to 60 cm.
6 | Working channel diameter should be of size 2.5 to 3.5 mm
7 | Outer diameter of the distal end should be of size from 6.0 mm to 6.5 mm.
8 | Insertion tube outer diameter should be of size from 6.0mm to 6.4mm.
9 | It should include the standard accessories as Carrying case, leakage tester, bite protector, cleaning brush, balloon applicator.
10 | Two or more no. of remote control switches on control body.
11 | Scope should be fiberless with CCD on tip , compatible with video processor.
B) | Pediatric Video Bronchoscope -01No
1 | Should have the facility of simultaneous suction and application of medication via removable suction valve.
2 | It should be high defination, waterproof, fully immersible for cleaning and Sterilizable.
3 | Tip deflection should be 180 up/ 130 down deg or more.
4 | Angle of view should be more than or equal to 110 degrees.
5 | Working length should be more than or equal to 60 cm.
6 | Working channel diameter should be of size 1.0 to 2.0 mm
8 | Insertion tube outer diameter should be of size from 3.5 mm to 4.5 mm.
9 | It should include the standard accessories as Carrying case, leakage tester, bite protector, cleaning brush, balloon applicator.
10 | Two or more no. of remote control switches on control body.
11 | Scope should be fiberless with CCD on tip , compatible with video processor.
C) | Ultrasound Linear Bronchoscope- 01 No
1 | Should have the facility of simultaneous suction and application of medication via removable suction valve.
2 | It should be high defination, waterproof, fully immersible for cleaning and Sterilizable.
3 | Tip deflection should be 120 UP/ 90 down deg or more.
4 | Angle of view should be more than or equal to 80 degree with a forward oblique view should not be more than 45 degree. Please specify the degree of forward view.
5 | Working length should be more than or equal to 60 cm.
6 | Working channel diameter should be of size 2.0 mm or more and it should be compatible for TBNA needle.
8 | Insertion tube outer diameter should be of size from 6.5 mm to 7.5 mm.
9 | depth of field of 3 mm-50mm or more.
10 | Ultrasound display modes should be B mode/ color mode.
11 | It should have scanning frequency range form 5MHz to 12 MHz or more.
12 | It should have scanning range 60 degree or more.
13 | Scanning direction should be longitudinal with scanning system- Convex
14 | It should be designated for both Balloon method & direct contact.
15 | It should include the standard accessories as Carrying case, leakage tester, bite protector, cleaning brush, balloon applicator.
16 | Two or more no of remote control body.
17 | Scope should be fiberless with CCD on tip , compatible with video processor & Ultrasound system.
D) | Digital Portable Ultrasound Scanner
1 | It should be compact & easily transportable unit. It should be ultrasound portable color Doppler Machine.
2 | Dual mode cpabilities to see in real time both B mode and CFM mode.
3 | Elastography facility should be available to visualize tissue strain in real time, offering information of its stiffness.
4 | high resolution imaging with HI COMPOUND TECHNOLOGY (with adjustable modes) to transmit and receive ultrasound beams at multiple angles and overlaps the images in real time which observes clear lesions through the reduction of speckle noise and improvement of contrast resolution.
5 | Ergonomic operation keyboard with color coded buttons for ease of operator use.
6 | It should have annotation, arrow mark and pointer display.
7 | USB should be available for external storage facility.
8 | Tissue Harmonic Imaging that uses broadband technology for better penetration and resolution should be avaial,ble for In-depth and good resolution in image quality. This technology basically expands the bandwidth of harmonic signals to maximum, making it possible ti include the low frequency subharmonic signals.
9 | Should be capable of giving multiple frequency range (5 MHz to 12 MHz) or more.
10 | It should have inbuilt/ external hard disk of atleast 1 TB
11 | It should have GAIN, Contrast, sTC functions.
12 | Cine Memory should be 150 frames or more.
13 | Should have DICOM and PACS compatibility.
C) | Spacs fo video Processor
1 | Compatible with Analog and Digital (Y/C, RGB, COMPOSITE, DVI) to reproduce high definition (1080p/i) images/ video.
2 | White balance adjustment should be availabel.
3 | Colour adjustment (Red, Green, Blue) should be available.
4 | It should have facility for image enhancement endoscope iscan/ NBI OR FIC technology for better diagnosis.
D) | Specs for Light Source.
1 | It should be a 300 W Xenon light source.
2 | Lamp life should be 500 hrs.
E | Atleast 26" Medical grade Monitor to be provided with complete system for visualisation.
F | Image management and recording software (PC based or built in) to supplied, along with necessary hardware with atleast 500 GB storage.
G | Should be supplied with crt/ trolley.
H | Accessories
1 | Balloon: 50Nos
2 | TBNA needle (EBUS): 50Nos
3 | EBUS transbronchial biopsy needle-50Nos
Should perform yearly calibration and preventive maintenance (2nos) during warranty & if entered into AMC(includes unlimited breakdown calls) /CMC (which will include spare replacement, consumable spares,breakdown calls) Testing & measuring equipment used should be traceable to SI units through National/international standards( As per NABH norms)
Please specify list & cost of consumables/consumable Spares( i.e.spares need to be replaced at regular intervals, maybe quarterly/half yearly/yearly such as annual maintenance kit etc.) if any.
Please specify pre installation requirement (electrical, HVAC etc)
Please specify fpootprint size & weight
Demo of the quoted model , will be required if so desired by the user, after the opening of the technical bid and prior to opening of financial bid. This is for technical evaluation.
Predispatch inspection at facory will be required if desired by authorities.
Plesae visit the site to check whether quoted model can be accomodated in the avaialble space.
Plesae specify power consumption.
International standards : The unit should comply with international standards and should have CE or US FDA marking or BIS marking or ICMED
Local Service Support: Should have local office and service support/service engineer for attending the breakdown calls.
Response Time: Should not be more than 6 hrs from lodging a breakdown complaint on toll free or by email.
Warranty: 2 years
CMC- CMC @ maximum 5% per annum for 8 years on Ex-works cost after warranty
Back to back assurance to be taken by the supplier from the OEM to supply spares for minimum 10 years and to be submitted.
1. If 2 years warranty, CMC @ maximum 5% per annum for 8 years on Ex-works cost after warranty and buyback conditions wherever applicable are not adhered to, TMC reserves right to reject the offer.
2. Please mention whether UPS/Stabilizer is required for the equipment. If so should be included in the offer.
3. Only Yes or No should be mentioned in compliance report and additional information should in remarks column.