## Multiparameter monitor
TATA MEMORIAL HOSPITAL
(PURCHASE DEPARTMENT)
Dept : | (Sangrur, Punjab)
Name of equipment :Multiparameter patient monitor
Please quote all the models including the state of the art equipment which meets the basic minimum specifications.
Sr. No. | Description of Item | Compliance
|  | Yes | No
Model Name :
Make :
1 | Should be portable, light weight not more than 4.5 kgs, patient monitor with basic parameters and additional capnography feature as option.
2 | Should be able to monitor ECG, respiration, SP02 NIBP, temperature and optional CO2.
3 | Should have high resolution color TFT display with size minimum of 8" diagonally.
4 | Should have optical encoder along with short cut keys for all important functions and should be user friendly.
5 | ECG:
a | Either 3 or 5 lead selectable .
b | Heart Rate range: should be between 20 -350 bpm.
c | Should have resolution upto+/-2bpm
d | Accuracy: lbpm or 1%, whichever is greater.
6 | Respiration:
a | Should be able to measure respiration rate from ECG.
b | Respiration range atleast of 4 to 50 BPM.
7 | NIBP:
a | Should measure with Automatic oscillometric method.
b | Manual/Automatic operation.
8 | Sp02:
a | Should have range of spo2 values: 1-100%
b | Accuracy: adult 100-70% (+2 digits), others please specify.
c | Neonates: 100-70% ( +3 digits)
d | Perfusion Indicator:- graph showing signal strength
e | PR Range should be atleast 20 to 230 BPM.
9 | Temperature:
a | Should be dual channel temperature measurement feature.
b | Range: 0°C to50°C.
c | Selectable unit like °C or °F.
10 | Capnography:
a | Should be either infrared side stream or micro stream capnography with display of C02 waveform and digital values of EtC02, FiC02 & RR
b | C02 gets priority for calculating respiration rate i. EtC02 Range: 0-99 MMHG
c | Warm up time should be less than 1 minute.
d | RR range from capnography should be 0-150 BPM
11 | Alarms
a | Should have Alarm Setting For All Parameters
b | Should have flashing with audio and message indicator
c | Should be able to set alarm level individually for HR / PR, SP02, RR, temperature & C02.
12 | Should have battery type either lead acid or Lithium Ion (Built in) with battery backup of minimum 2 to 4 hours
13 | Battery charging time should be not more than 8 hrs (typically).
14 | Warranty : 2 years
15 | AMC : 5 Years free
PURCHASE OFFICER